SOP HISTORICAL RECORD

| DATE | WRITTEN/REVISED* BY | SUPERVISOR REVIEW/APPROVAL | MD REVIEW/APPROVAL |
| --- | --- | --- | --- |
| 10-5-21 | E. Skelly | BMH 10/5/21 | RK 10/5/21 |
| 10/5/2022 | E. Skelly | BMH 10/6/22 | RK 10/6/22 |

| *REVISION DATE | REVISION DESCRIPTION | ALL COPIES REVISED? |
| --- | --- | --- |
| 10/5/22 | Updated to include Ambulatory Surgery (sequence #5) | yes |

DOCUMENT CONTROL

| EFFECTIVE DATE: | 10-5-21 |  |
| --- | --- | --- |
| LOCATIONS OF ALL COPIES: | [MASTER - electronic] Lab-Dept. Shares |  |
|  | Ancillary manual-lab |  |
|  | Ancillary manual-CLC |  |
| REMOVED FROM SERVICE: |  |  |

Purpose

This procedure will outline the steps required to handle proficiency 				samples, record results, submit these results to the College of 				American Pathologists, and how to process evaluations once they are 			received.

Policy

CAP surveys are sent two times per year to each area performing BinaxNOW point of care testing. These survey samples are to be tested the same as 	patient samples. Results are to be submitted directly to CAP online via the CAP website from each testing area.

Sample

CAP Proficiency test SARS-CoV-2 antigen, three specimens. Store in the refrigerator at 2-8C until testing. Once open, start testing immediately. Sample analysis and submission of results must be completed within the specified time limit.

Supplies

Test Kit: Abbott BinaxNOW Covid-19 Ag Cards– storage (2-30 degrees C) until the expiration date on the label.  Do not open the pouch until ready to perform the assay.

Special Safety Precautions

- Personal Protective Equipment (gloves) should be utilized.

- Disinfect work surfaces after testing.

- All survey samples should be treated as if potentially infectious and should be handled as if they are capable of transmitting disease.

Procedure

| STEP | ACTION |
| --- | --- |
|  | Person who receives the survey in the lab is to fill out the initial routing slip “survey received for testing” (Attachment A). This is to stay with the samples and paperwork until completion. |
|  | Proficiency testing will be performed by individuals that routinely perform testing. The samples will be rotated among the staff as a part of competency assessment. |
|  | Thoroughly read the kit instructions before proceeding. (Attachment B) |
|  | Following CAP instructions for BinaxNOW testing by 1st inserting a swab from the BinaxNOW kit into the proficiency sample provided by CAP for 15 seconds until thoroughly saturated. Then test covid-19 antigen proficiency samples following SOP AT14 in the same manner as patient testing, including filling out the BinaxNOW Covid-19 Ag Card Result log sheet. |
|  | There should be NO communication regarding the results of proficiency testing samples with other laboratories or locations. The samples should not be sent to another laboratory for analysis before submitting the result to the proficiency testing provider. |
| 6. | When testing is completed, fill out the result sheet under the appropriate specimen number.(Attachment C) Make sure that the correct method is listed (1872) All testing personnel should sign the attestation form and then give to the nurse manager to review for legibility and completeness. The nurse manager will sign as the “director or designee”. Be sure to fill out the survey line next to the director signature. |
| 7. | The original PT answer sheet, and the attestation form are submitted online via the CAP website (www.cap.org) (Attachement D) to the proficiency testing program provider within the specified time periods by the nurse manager or designee. RESULTS MUST BE RECEIVED BY CAP BEFORE THE DEADLINE. RESULTS CANNOT BE ACCEPTED IF RECEIVED AFTER THE DUE DATE. The website confirmation pages will be saved with the copies of submitted results. Print the online screen from the “RESULT FORM DATA ENTRY” where the status shows “Received”. |
| 8. | Send the samples down to the lab after testing for storage. The material maybe needed for corrective action if performance was not acceptable. |
| 9. | Once the survey results have been evaluated and graded, the Ancillary Testing Coordinator will send signed copies to testing site (signed by Ancillary Testing Coordinator, P&LM Supervisor, and Medical Director). Results will be reviewed by the persons who performed the testing, and the person in charge. The nurse manager should also sign the result evaluation. |
| 10. | Once step 9 is completed, file the result evaluation forms in CAP binder for that year with the results submitted. |

Expected Values

Performance Criteria for a single proficiency testing event are as follows:

#### Satisfactory:

- A score of 80% or greater for analyte(s) which is composed of five samples, a score of greater than 50% for any analyte in an event less than 5 challenges.

- Returning proficiency testing results to the proficiency-testing provider within the timeframe specified by that provider.

- No omission of results on the proficiency testing form.

- Participation in a proficiency-testing event.

- Unsatisfactory:

- Failure to attain a score of at least 80% for analyte(s) in all specialties, subspecialties or tests with 5 challenges per survey, 50 % or less for any analyte composed of less than 5 challenges.

- Failure to return proficiency testing results to the proficiency testing provider within the time frame specified by the provider resulting in a score of zero (0).

- Omission of results or method/instrument codes on the proficiency testing form.

- Failure to participate in the proficiency testing event.

Result Interpretation

Any score of less than 100% will be reviewed and remedial action 		documented.   The Ancillary Testing Coordinator will review the procedure, 	conduct a comprehensive investigation to determine the problem, evaluate 	patient results at the time of the unsatisfactory PT performance, and take 	appropriate corrective action.

Unsuccessful Performance as defined by CLIA and Joint Commission:

Unsuccessful performance is defined as a failure to achieve satisfactory performance for two (2) consecutive or two (2) out of three (3) consecutive testing events for any analyte tested.

In the event of unsuccessful performance, the Ancillary Testing Coordinator will           follow established laboratory policy and launch further investigation.

References

- The Joint Commission Accreditation Manual for Laboratory & Point-of-Care Testing, current edition.

- VHA Handbook 1106.1, Enforcement Requirements, current edition.

- “How to Meet Most Frequently Cited Laboratory Standards”, JCAHO Resources, 2001.

Attachments

- Routing slip: Survey received for testing

- Example: CAP kit instructions

- Example: blank CAP result sheets

- Entering Results - CAP Website