Meridian Health

Jersey Shore University Medical Center

### Laboratory

Neptune, New Jersey 07753

| SUBJECT: PARAFFIN MICROTOMY | TITLE: PARAFFIN MICROTOMY |
| --- | --- |
| EFFECTIVE DATE: 4/2013 | SECTION: HISTOLOGY |
| APPROVED: | POSITION RESPONSIBLE FOR REVIEW: Laboratory Director or Designee |

PARAFFIN MICROTOMY

Principle: To consistently produce sectioned thickness of paraffin embedded tissue for various

tissue types and procedures.

Procedure:

- Assess type of tissue.

- Check microtome settings and adjust appropriately depending on tissue or procedure. Use guideline from table (below).

- Shave block to remove paraffin and expose tissue.

- Gently rotate microtome to produce a “ribbon.”

- Use “hot” & “cold” technique if necessary.

- Place “ribbon” on the waterbath.

- Pick up “ribbon” with slide (no holes and/or folds).

- Take block and copy accession number on slide/s.

| Tissue/Procedure | Thickness |
| --- | --- |
| Routine Paraffin | 4 to5 microns |
| Biopsies (Small) | 4 to5 microns |
| Bone Marrow (clot) | 2 to3 microns |
| Bone Marrow biopsy | 4 to 5 microns |
| Amyloid Stain | 6 to12 microns |
| Lymph Node | 3 to 4 microns |
| IHC Stain | 3 to 4 microns (use positively charged slides) |
| Special Stains (Fe, Retic, Tric, Mucin, GMS) | 4 to 5 microns |
| Frozen Sections | 4 to 6 microns |
| Femoral Heads | 5 to 6 microns |

| Revised: | Reviewed: | Reviewed | Reviewed: |
| --- | --- | --- | --- |
| Reviewed: | Reviewed: | Reviewed: | Reviewed: |
| Reviewed: | Reviewed: | Reviewed: | Reviewed: |
| Reviewed: | Reviewed: | Reviewed: | Reviewed: |