| # PURPOSE | This procedure describes how to do an automated body fluid cell count using the Sysmex XN-9000 hematology analyzer. |
| --- | --- |

| # POLICY | Quality control and limitations of the procedure must be taken into account before validating automated body fluid cell count results into the laboratory information system (LIS). |
| --- | --- |

| # SCOPE | This procedure is intended for the use of Clinical Lab Scientists (CLS) that will use the XN-9000 to do an automated body fluid cell count and other laboratory personnel who may need to review the assay as part of quality management. |
| --- | --- |

| # REAGENTS | XN-9000 Reagents Anticoagulants/Additives: - EDTA (Other anticoagulants may alter test results). |
| --- | --- |

| # EQUIPMENT / SUPPLIES | XN-9000 Calibrated MLA Pipettes for dilutions |
| --- | --- |

| # SAMPLE/ # CAUSES FOR REJECTION | Acceptable Body Fluid Types are: - Cerebrospinal fluid (CSF) - Serous fluids (Peritoneal, Pleural, Pericardial, etc) - Synovial Fluids with added hyaluronidase to break up mucous. Required sample volume 1.0mL or more Aspirated sample volume Approximately 88 μL Clotted specimens will not be run automated due to the mucous material that could clog up the instrument. |
| --- | --- |

| # SPECIAL SAFETY PRECAUTIONS | Whenever working with body fluids, universal precautions must be followed. Proper protective equipment and wear must be used. Refer to the safety manual for general safety requirements. |
| --- | --- |

| # QUALITY CONTROL | XN CHECK BF - Automated Body Fluid Controls Level 1 Level 2 |
| --- | --- |
|  | Both levels of XN CHECK BF controls must be run daily per XN instrument per shift, unless instrument is out of order. Graveyard shift will only run XN CHECK BF controls on one analyzer and only if a body fluid analysis is performed during their shift. |

PROCEDURE:

| # Analysis Preparation | # Step | # Action |
| --- | --- | --- |
|  | # 1. | Check the status of the analyzer. Confirm the analyzer is ready. |
| # Background Check | # 2. | Press the mode switch to eject the tube holder. |
|  | # 3. | Select the Change Analysis Mode button on the control menu. |
|  | # 4. | Select BODY FLUID |
|  | # 5. | Analyzer will automatically perform a Background Check |
|  | # 6. | Ensure Background Check passes, then proceed to next step. |

| Body Fluid Analysis | Step | Action |  |
| --- | --- | --- | --- |
| # Control Analysis | 1. | Bring the XN CHECK BF to room temperature for 15 minutes |  |
|  |  | Mix XN CHECK by gentle end to end inversion (approximately 20 times or until the cell button is suspended) |  |
|  |  | Place a well-mixed vial of XN CHECK BF in the sample tube holder |  |
|  |  | Click the Manual Analysis button in the analyzer area |  |
|  |  | Confirm that READ ID is checked |  |
|  |  | Confirm that CAP OPEN is unchecked |  |
|  |  | Click OK |  |
|  |  | Press the start switch (blue button) |  |
|  |  | Verify that controls are within acceptable range. If not, rerun. If still not within expected range, do not use instrument for automated body fluid cell count. If acceptable, proceed to sample assay. |  |
|  | 1. | Place well-mixed patient body fluid in a vial with sample barcode for analysis and place in the sample tube holder. |  |
|  | 2. | Click the Manual Analysis button in the analyzer area. |  |
|  | 3. | Confirm that READ ID is checked. |  |
|  | 4. | If sample tube is uncapped, click CAP OPEN. If sample tube is capped ensure the CAP OPEN box is unchecked. |  |
|  | 5. | Click OK and Press the start switch (blue button). |  |
|  | 6. | Print hard copy of result. Review for analysis errors (*) or linearity flags (@) and correct. |  |
|  | 7. | Perform calculations and write your results on the printout. |  |
|  | 8. | Have 2nd CLS review result for accuracy and initial printout. |  |
|  | 9. | Report in Cerner ARE |  |

______________________________________________________________________

| # CALCULA-TIONS | RBC-BF results must be multiplied by 106 (Example: RBC-BF = .005 X 106 Move decimal six places to the right. Result should be reported as 5000) TC-BF results must be multiplied by 103 (Example: TC-BF = .055 X 103 Move decimal three places to the right. Result should be reported as 55) |
| --- | --- |

| # INTERPRETATION/ RESULTS | Report from printout the RBC-BF result and the TC-BF result. RBC-BF is resulted in Cerner as RBC Auto. TC-BF is resulted in Cerner as TNC Auto. Reportable Range: TC-BF: 0.003 – 10.000 x 103/uL or 3 – 10,000 RBC-BF: 0.002 – 5.00 x 106/uL or 2,000 – 5,000,000 For results over reportable range, dilute with Cellpack DCL to bring result within range,( ie TNC = 50,000@ – Dilute 1:10, rerun, multiply result X10). For results less than reportable range, perform manual count. |
| --- | --- |

| # REFERENCE INTERVALS | See Body Fluid Cell Count Procedure |
| --- | --- |

# REFERENCES: 	Sysmex XN-9000 Instructions for Use (North American Edition), Sysmex Corporation, Kobe, Japan.

# RELATED DOCUMENTS: Body Fluid Cell Count Procedure

# AUTHOR:	Nanette Taft, CLS, MT(ASCP)

Error! Cannot open data source.DOCUMENT HISTORY PAGE

Effective Date: 10/18/19

| Change type: new, major, minor etc. | Changes Made to Document – Describe | Area Lab Manager Reviewed/ Date | Lab Operations Director Reviewed/ Date | Medical Director Reviewed/ Date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| NEW |  |  |  |  |  |