#### Safety Message     Use appropriate universal precautions when performing the test.

| ##### Introduction | The Sysmex XbarM quality control program collects batches of reliable (not necessarily normal) patient results and applies Dr. Brian Bull’s formula to all the primary parameters as well as the sensitivity parameters. The algorithm produces a “weighted average” of patient data and is plotted in a QC chart. Standard statistical process control allows the stability of patient based averages to be monitored and flagged. XbarM QC compliments the use of XN Check control material. Where control drift or matrix effect is an issue, XbarM may aid in determining if there are real quality issues by evaluating the effect on patient results. |
| --- | --- |

| ##### Policy | XbarM is an internal quality control process used in conjunction with Sysmex XN Check QC to monitor the performance of the Sysmex XN 9000 Hematology Analyzers. Action is required when results fall outside the control limits. |
| --- | --- |

| Summary | The formula and calculations are incorporated into the computerized data analysis. This laboratory will establish XbarM (Xm) historical limit % using 200 data points, representing 4000 samples in 20 patient size batches. Historical Xm limits are collected over multiple reagent lots and at least one (1) month, including all types of patient samples normally encountered. Batch size for Xm will be 20 patient samples per batch. Each point on the Xm QC graph represents one batch. Each shift will review Xm charts when they run their daily QC. Manager will review Xm charts at least monthly. |
| --- | --- |

| ##### Establishing the X-barM Control Program ##### Monitoring XbarM graphs | Step; Action. 1.; Manually enter the Suggested File Limits for X-barM while collecting baseline data.. 2.; Use the variableTarget function during data collection. - Click on the QC icon from the menu window or Toolbar. - Click on each X-barM file. - Click Target/Limit - Click Variable Target, and the OK. 3.; Auto set targets, and establish a 3 CV% limit for all parameters that will be monitored once at least 200 data points are collected over multiple reagent lots and at least 1 month. 3CV% limits provides a reasonable level of error detection for most parameters.. Follow these steps to monitor XbarM graphs Step; Action. 1.; Click on each XbarM file for each analyzer in the QC module of the XN 9000 IPU.. 2.; Start with the CBC file and scroll through all parameters.. 3.; The calculated values are shown as a red cross if they exceed or fall below the limits set.. 4.; Action is required for these values that fall outside the control limits after two consecutive red crosses. Record in action log.. 5.; Be sure to check the Diff, Retic and PLT-F files also.. 6.; Initial on Daily Hematology Checklist after all parameters have been monitored each shift.. |
| --- | --- |
| ##### New Reagent Lot Verification | Performance verification of new reagent lots and/or shipments may be accomplished by monitoring XbarM of reportable parameters. Step; Action. 1.; New lot of reagent installed on instrument.. 2.; Review XbarM after the next two (2) batches of 20 samples. XbarM is the most sensitive method for detecting changes with a new lot of reagent.. |

| ##### Trouble- Shooting results | Follow the steps below if the XbarM is out of limits. Step; Action. 1.; A red cross will flag on the Xm chart.. 2.; Check to make sure samples were run randomly and batches were not predominately Peds or Oncology patients. If samples were not a random population then document this in action log.. 3.; If XbarM continues to exceed limits with 2 or more red crosses, then check instrument calibration and precision by running one level (L1, L2 or L3) of XN Check QC.. 4.; Visually inspect reagents to ensure they are installed correctly and to see if a new lot of reagent has recently been installed. Performance verification of new reagent lots and/or shipments may be accomplished by monitoring Xm of reportable parameters.. 5.; Investigate all unusual results. Resolve problems and document.. 6.; Notify Hematology Manager, Sysmex Technical Assistance Hotline or Regional Laboratory Instrumentation if unable to correct problem.. |
| --- | --- |

| ##### Author | Nanette Taft, CLS (ASCP) |
| --- | --- |

| ##### Derivative Documents | Sysmex, XN – Series Resource Manual Document Number: 1002-LSS, Rev. 2 – February 2013 |
| --- | --- |

| ##### Distribution | South Bay Laboratory |
| --- | --- |

DOCUMENT HISTORY PAGE

Effective Date: 05/14/18

| Change type: new, major, minor etc. | Changes Made to Document – Describe | Area Lab Manager/ Date | Lab Director Reviewed/ Date | Med.Dir reviewed/ Date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| NEW |  |  |  |  |  |