| ##### Policy | Quality control files for the Sysmex XN Check and XN Check BF Controls are automatically uploaded after each run via SNCS connection to the Sysmex Insight Quality Assessment Program. The data is used for evaluation and comparison with other XN9000 analyzers. The Interlaboratory Quality Assurance Program (IQAP) is a service provided by Sysmex Corporation. The IQAP report provides interlaboratory comparison indicating precision and accuracy relative to peer data. |
| --- | --- |

| ##### Print Control File: | Follow the steps below for printing the control files after each period ends: |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | On IPU, click on QC files icon. Double click on the file to print. QC chart is displayed. |
| 2. | Set Range of QC data to print. Click on Range and capture the points with the cursors. |
| 3. | Select Output to print the selected chart to both GP and LP. |
| 4. | Print Levy-Jennings charts (GP) and data (LP). |
| 5. | Repeat steps 1-4 until all files are printed. |
| 6. | Insight reports are generated and reviewed at least monthly by Hematology Manager. |

| ##### Reviewing the IQAP Report: | IQAP data report can be generated immediately after upload, however, all peer groups may not have been received. The IQAP program determines how the laboratory instrument compares with participant instrument peers. The Hematology Manager reviews the IQAP report for accuracy and precision, investigates possibility of system error and documents results of investigation on the IQAP report. |  |
| --- | --- | --- |
|  |  | Mean Diff This is an expression of the difference of your analyzer’s mean from the peer group. It can be used to determine the absolute magnitude of an accuracy error. |
|  |  | SDI This is a standard measure of your analyzer’s inaccuracy. It provides the distance of your mean from the “truth” (the peer group mean) in units of standard deviations. Positive and negative SDIs indicate error above and below the Group Mean, respectively. Zero values may occur when Your Mean is close to Group Mean. The formula is the same as used for “z-score” or “SDI”. |

|  |  | Your CV This ratio is a measure of your precision, which can be compared across control levels and lot numbers. |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Reviewing the IQAP Report: |  | (continued) Group CV This is a measure of the level of precision obtained by the laboratories contributing to the peer group statistics. It does not reflect the average obtained by the group, but is expressed as the degree of variability within the group in a manner consistent with that for assessing the individual laboratory. If Your CV is greater than 1.5 times the Group CV, your result is presented in bold type. Note: When Your CV result is greater than the Group CV, this does not necessarily indicate that your precision is unacceptable. Both CV and SD should be reviewed to assess analyzer performance. CV’s of extreme levels are usually due to anomalous data. |  |  |  |
|  |  | IQAP Checklist: |  |  |  |
|  |  |  | Situation | Action |  |
|  |  |  | Your data agrees with that of other participating labs: no flags | No action necessary. |  |
|  |  |  | Your SDI for one or more parameters/levels is flagged with an accuracy bias code. (See attachment A). | Your accuracy needs to be reviewed. - Verify that the data on the report matches the data you submitted and that your data was submitted correctly. - Check for a pattern by comparing the data with other levels of the same parameter on the report. - Check for similar patterns in both previous and current lots of control. - Document review on report. |  |

| Reviewing the IQAP Report: | (continued) |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | IQAP Checklist: (continued) |  |  |  |
|  |  |  | Situation | Action |  |
|  |  |  | Your CV for one of more parameters/levels shows a significant change from previous month or is highlighted BOLD on report. | Your precision needs to be reviewed. - Verify that the data on the report matches your records and that you submitted the data correctly. - Review maintenance procedures. - Review handling techniques of the cell control product. - Document review on report. |  |
|  |  |  | Both SDI and CV for one or more parameters/levels should be reviewed. | First resolve the precision problem, then troubleshoot the accuracy problem. Follow the steps listed above. Contact Sysmex TAC Hotline for assistance |  |

| Author | Nanette Taft, CLS, MT(ASCP) |
| --- | --- |

| Attachments | A | Peer Group Comparison |
| --- | --- | --- |

| ##### References | - Insight ™ Participant Report Guide, Document Number: 1090-TS, Rev.4 9/2015 |
| --- | --- |

| Related Documents | Documents | Title |
| --- | --- | --- |
|  | LHS 2510 | Sysmex XN 9000 |

| ##### Distribution | - South Bay Laboratory |
| --- | --- |

Document History Page

Effective Date:  05/14/18

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Area Lab Manager/date | Lab Director Reviewed/ Date | Med. Dir. Reviewed/ Date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| NEW |  |  |  |  |  |

Attachment A

A. Peer Group Comparison

The second page of the Period Report displays the Peer Group Comparison for each

parameter. This section provides comparison of the laboratory’s performance with regard to

accuracy and precision. An example is shown below.