| ##### Safety ##### Message | All Quality Control reagent products contain biologically source materials from human, avian, reptile and ungulate. Always use appropriate universal precautions while using them to perform a test. |
| --- | --- |

| ##### Policy | - All unacceptable control results are resolved before testing patient specimens. - QC only run when instrument is in use or test is ordered on shift. - When controls are out of range on blood cell analyzers, a step by step approach must be taken to resolve the out of control results on commercial controls. |
| --- | --- |

| ##### Definition of out of control | "Acceptable limit of error" - the permissible range by which the obtained results may vary from the specified result. Repeated analysis of the control or an unknown specimen should give results within this limit. "Out-of-Control" - If a result outside of these limits is obtained. |
| --- | --- |

| ##### Interpreting results and course of action: | If values are outside the acceptable range, the data must be carefully evaluated before any results are reported. If the source of error cannot be found, a new control specimen along with several of the unknowns chosen at random is re-assayed. If the new control gives an appropriate value and the unknowns check with the previous tests performed the same day, the values are reported; if not the data is held back and the reagents procedure, and instrument are checked one by one until the problem is found and resolved. |
| --- | --- |

| ##### Equipment ##### And ##### Reagents | ### Sysmex XN 9000 - Sysmex XN Check tri-level control - Sysmex XN Check BF bi-level control - iSED Seditrol bi-level control - Alcor iSED - OSOM® Mono Test Kit - Oraquick Rapid HIV Test Kit - MES SQA-Vision - MES Qwik Check Beads - Rapid Malaria Kit/ Positive Malaria Control - Incyto disposable Hemacytometer - Streck Cell-Chek |
| --- | --- |

| Control Run Outline Table: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Control | Instrument | Frequency | Performed by: | Reviewed |
| Oraquick Rapid HIV | Rapid HIV | - Daily with each new operator when test is ordered. | Clinical Lab Scientist | At least monthly by a supervisor. |
| MES Qwik Check Beads | Vision | - When analyzer is in use, once every 24 hours. | Clinical Lab Scientist | At least monthly by a supervisor. |
| Mono Test Control | Mono Test Kit | - External controls should be run once daily for each day of use.. | Clinical Lab Scientist | At least monthly by a supervisor. |
| Cellavision QC slide | DI-60 | - When analyzer is in use, once every 24 hours. | Clinical Lab Scientist | At least monthly by a supervisor. |
| XN Check Controls – L1, L2, L3 XN Check BF Controls L1, L2 | XN-9000 | - When analyzers are in use, three levels of XN-Check and two levels of XN Check BF controls are run once per shift for each instrument. Night shift will only run XN Check BF on one XN when a body fluid cell count is ordered. | Clinical Lab Scientist | At least monthly by a supervisor. |
| XM | XN-9000 | - Monitor results each shift . | Clinical Lab Scientist | At least monthly by a supervisor. |
| Instrument Correlation | XN-9000 iSED | - Biannual evaluation is performed. | Clinical Lab Scientist | Biannually by supervisor. |
| Alcor Seditrol I &II | iSED | - Daily by AM shift | Clinical Lab Scientist | At least monthly by a supervisor. |
| Rapid Malaria Positve Control | Rapid Malaria Kit | - When kit is in use, once every 24 hours. | Clinical Lab Scientist | At least monthly by a supervisor. |
| Streck Cell Chex | Hemacy-tometer | - AM, PM shift daily. - Night shift only when test is ordered. | Clinical Lab Scientist | At least monthly by a supervisor. |

| Procedure |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | I. DAILY |  |  |  |  |
|  | AM Shift: |  |  |  |  |
|  |  | Control |  | Instrument or Equipment |  |
|  |  | XN Check (Level 1,2,3) XN Check BF (Level 1,2) LJ/XbarM charts |  | XN-9000-1, 2, 3 |  |
|  |  | CellaVision QC slide |  | DI-60 |  |
|  |  | Alcor Seditrol I & II |  | ISED |  |
|  |  | Streck Cell-Chex: Alternate L1 & 2 |  | Hemacytometer |  |
|  | PM Shift Daily QC: |  |  |  |  |
|  |  | Control |  | Instrument |  |
|  |  | XN Check (Level 1,2,3) XN Check BF (Level 1,2) LJ/XbarM charts |  | XN-9000-1, 2, 3 |  |
|  |  | Streck Cell Chex: Alternate L1 & 2 |  | Hemacytometer |  |
|  | NIGHT Shift Daily QC: |  |  |  |  |
|  |  | Control |  | Instrument or Equipment |  |
|  |  | XN Check (Level 1,2,3) XN Check BF (Level 1,2) LJ/XbarM charts |  | XN-9000-1, 2, 3 XN-9000 1,2 or 3 (For BF) |  |
|  |  | Streck Cell-Chex Alternate L1 & 2 |  | Hemacytometer |  |
|  | - Note: Cell count control and XN Check BF is only run when fluid cell count order is received on night shift. - Mono Test Controls are run once for each day of use. |  |  |  |  |
|  | - a. |  | - Record reagent lot #s and expiration dates in QC log. |  |  |
|  | - b. |  | - Perform test using the control material. |  |  |
|  | - c. |  | - Record results in designated area of QC log or instrument data base. |  |  |
|  | - d. |  | - Evaluate control results and resolve out of control runs before proceeding with patient testing. (see section III of this procedure) |  |  |

|  | - II. UNSCHEDULED: |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | A. Parallel Testing of New Control Lot #: |  |  |  |  |
|  |  | Control | Instrument or equipment |  |  |
|  |  | Mono Test | Mono Test Kit | Test each new lot or shipment with current controls |  |
|  |  | Rapid HIV | Rapid HIV Kit |  |  |
|  |  | Rapid Malaria | Rapid Malaria Kit |  |  |
|  |  | Qwik Check Beads | Vision |  |  |
|  |  | XN Check, XN Check BF | XN-9000 | at least 5-10 runs to establish/verify range |  |
|  |  | Alcor Seditrol | iSED |  |  |
|  |  | Streck Cell Chex | Hemacytometer |  |  |
|  |  | Note: Ideally to establish/verify new lot number control range, parallel testing is conducted over several days. i.e. 5-10 days. If the new control material arrives late, it may be necessary to perform an abbreviated parallel testing run. For the Alcor iSED ESR Control and the Streck Cell Chex controls a verification run is performed. Verification run must be within assay range of insert. When enough data is collected, mean and 2 SD is determined and new range is applied. |  |  |  |

|  | III. Evaluating Control Run: |
| --- | --- |

|  |  |  |  | - Run Controls |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | - NO |  |  | - Within Acceptable Limits |  |  |  |  |  |  |  |  | - YES |  |  |
|  | - Document results in corrective action log. Circle and initial values. Check reagents for, volume, outdate (age/expiration date), correct lot number, and signs of contamination. Run controls again. |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. |  |  |  |  |  |
|  | - NO |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  | - YES |
|  | - Document results in corrective action log. Open a new vial of control. Run controls again |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. |  |  |  |  |  |
|  | - NO |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  | - YES |
|  | - Document results in corrective action log. Troubleshoot instrument. Document on corrective action log. Run controls again. |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. Discard bad reagent. |  |  |  |  |  |
|  | - NO |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  | - YES |
|  | - Document results in corrective action log. Get assistance from Regional Instrument Repair department. Notify supervisor. Use back-up instruments. Run controls on backup instruments if needed. Document results. |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. |  |  |  |  |  |
|  | - NO | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  |  | - YES |
|  | - Notify supervisor, and floors. Call instrument repair (8-397-7000) |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. |  |  |  |  |  |

| ##### Troubleshooting Controls |  | Use the following table to determine who is responsible for evaluating quality control results and correcting problems. |  |  |
| --- | --- | --- | --- | --- |
| Stage | What Happens |  | Who is responsible |  |
| XN Check Control | Evaluate the results of the XN Check control. If the results are not within the assigned values, troubleshoot and document the steps necessary to resolve the problem. |  | Clinical Laboratory Scientist |  |
| XM | Evaluate the XM batches and if the results are drifting troubleshoot the analyzer and document the steps taken to resolve the problem. |  | Clinical Laboratory Scientist |  |
| Instrument Correlation | Evaluate results. If results are not within the designated limits, troubleshoot and document the steps taken to resolve the problem. |  | Clinical Laboratory Scientist |  |
| Mono Test Controls | Evaluate results. If results are not within the designated limits, troubleshoot and document the steps taken to resolve the problem. |  | Clinical Laboratory Scientist |  |
| MES Quik Chek Controls | Evaluate results. If results are not within the designated limits, troubleshoot and document the steps taken to resolve the problem. |  | Clinical Laboratory Scientist |  |
| Rapid HIV Controls | Evaluate results. If results are not within the designated limits, troubleshoot and document the steps taken to resolve the problem. |  | Clinical Laboratory Scientist |  |
| Rapid Malaria Controls | Evaluate results. If results are not within the designated limits, troubleshoot and document the steps taken to resolve the problem. |  | Clinical Laboratory Scientist |  |
| Alcor iSED Seditrol Controls | Evaluate results. If results are not within the designated limits, troubleshoot and document the steps taken to resolve the problem. |  | Clinical Laboratory Scientist |  |
| Streck Cell-Chex I or II | Evaluate results. If results are not within the designated limits, troubleshoot and document the steps taken to resolve the problem. |  | Clinical Laboratory Scientist |  |

| - Procedural Notes | - All control testing is performed according to specific test procedure. |
| --- | --- |
|  | - Cell count control level run on shift is rotated between AM & PM shift. On the night shift, cell count control is only run when a body fluid cell count specimen order is received on shift. The control level to be run on the night shift is determined by which level was run the last time a test order was received on shift. e.g. If level I was run two days ago when a CSF cell count was performed on night shift the next time a cell count is performed on shift, level II must be run with test. - Mono Test controls are run once for each day of use. - Rapid HIV controls are run daily with each new operator when test is ordered. |

| ##### Author | Carlo E. Punu, MT, CLS Nanette Taft, MT (ASCP), CLS Mina Acosta, MT (ASCP), CLS |
| --- | --- |

| ##### References | College of American Pathologists: Laboratory Accreditation Manual 2007 July Edition by Francis Sharkey, MD, FCAP |
| --- | --- |

| ##### Related Documents | - Body Fluid Analysis Procedure - Body Fluid Cell Count Quality Control - CellaVision Procedure - Sysmex XN 9000 Procedure - ISED Procedure - MES Vision Procedure - Rapid HIV Procedure - Rapid Malaria Procedure - Mono Test Procedure - Hematology Instrument Correlation Procedure |
| --- | --- |

DOCUMENT HISTORY PAGE

Effective Date: 12/01/02

| Change type: new, major, minor etc. | Changes Made to Document – Describe | Signature responsible person/Date | Lab Manager Reviewed/ Date | Med. Dir. Reviewed/ Date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| NEW | Change in instrument technology | M. Acosta 10/04/04 | N. Muneno 10/07/04 | R.Doshi,MD 10/12/04 | 10/12/04 |
| MAJOR | Change frequency of 5C and Retic- C Controls runs per shift on both LH750 instruments. | Carlo Puno 02/16/09 | N. Muneno 02/19/09 | R.Doshi,MD 02/17/09 | 02/17/09 |
| MAJOR | Change in instrument technology to Sysmex. Addition of Mono Test. | N.Taft 09/15/11 | N.Muneno 09/16/11 | R.Doshi,MD 09/16/11 | 09/16/11 |
| MAJOR | Addition of Rapid HIV, Spermalite | N.Taft 08/24/12 | N.Muneno 8/28/12 | S.Wirio,MD 8/30/12 | 8/30/12 |
| MAJOR | Replacement of XE 2100 with new XN 9000, replacement of SureVu Color Mono with OSOM Mono Test. |  |  |  |  |