| Policy | - Use of a hematology analyzer to perform a complete blood count (CBC) and a leukocyte differential count occasionally requires a blood smear review or manual differential count to ensure accurate results. Consistent and clinically relevant criteria for smear review should be applied in order to maximize quality of results while preventing usage of time and labor to perform unnecessary slide reviews. - This regional policy outlines the criteria for smear review and manual differential count when an automated complete blood count (CBC) with or without automated differential count is ordered. - Peripheral blood smears are saved for 7 days. |
| --- | --- |

| - Specimen | - Acceptable Specimen - EDTA anticoagulated whole blood - Minimum Volume Requirement - 1.0 mL - Specimen Stability - Refrigerated (2-6oC) up to 24 hours. |
| --- | --- |

| - Equipment - And - Reagents | ### Sysmex XN9000 - CellaVision DI60 - SP 10 Slide Stainer - glass microscope slide - Hematek stainer - microscope |
| --- | --- |

- I-Slide Review/ Manual Differential Screening Procedure:

| Step | Action |
| --- | --- |
| - 1. | When to scan: CBC without auto diff. Parameter; Value or Flag(s). WBC; <2.0 thou/ul or >30.0 thou/ul. MCV; <65.0 fl or >120.0 fl. RDW; >22.0. Platelets; <20,000/ul (Adult)<30,000 (Peds) or >1,000,000/ul. Other Criteria; OP Alerts pertaining to Hemogram parameters.. |
| 2. | When to scan: CBC with auto diff Parameter; Value or Flag(s). WBC; <2.0 thou/ul or >30.0 thou/ul. MCV; <65.0 fl or >120.0 fl. RDW; >22.0. Platelets; <20,000/ul (Adult)<30,000 (Peds) or >1,000,000/ul. Neutrophils Absolute; <0.5 thou/ul or >20.0 thou/ul. Eosinophils Absolute; >2.0 thou/ul. Basophils Absolute; >0.5 thou/ul. Lymphocytes Absolute; >5.0 thou/ul. Monocytes Absolute when WBC <10,000.; >1.5 thou/ul for > 12 yrs. ;>3.0 thou/ul for ≤ 12 yrs. on first time for 2 week interval.. NRBC; ≥10%. Other criteria; Instrument error or flag prevents complete reporting of automated differential count.. Perform a Manual Diff when; Scanning shows abnormal cells or immature myeloid (more than 9 bands or more than 1 metamyelocyte)or any myelocyte, promyelocyte or any blasts.. |
| 3. | A differential count should not be performed on orders for CBC without an automated or manual differential count ordered. A physician order is required to perform testing not originally requested. If critical abnormalities are encountered during a scan without diff order, contact provider for order. |
| 4. | If the provider requires a manual differential to be performed in addition to a CBC (regardless of CBC results), the provider should order a manual differential in addition to a CBC without auto differential. |
| 5. | CLS must follow Regional WAM rules when reviewing slides. WAM rules that are triggered during analysis will appear in the OP alerts section of the WAM validation screen. |
| . |  |

- II-Preparing A Manual Peripheral Blood Smear Procedure:

| Step | Action |
| --- | --- |
| - 1. | Mix blood by inversion at least 10X. Insert DIFF-SAFE dispenser through the stopper of a tube held in the upright position. |

| Step | Action |
| --- | --- |
| - 2. | Turn the tube upside down and press moderately against a microscope slide, about ½ inch from one end of slide. The instant the drop appears, discontinue applying pressure. Note: If the drop fails to appear, check specimen for clots. For microtainer samples, make smear using wooden sticks. |
| - 3. | Lay the slide on a flat surface. Place the end of a second slide, held with the fingertips at an angle no greater than 300, in front of the drop of blood. |
| - 4. | Pull the spreader slide back into the drop of blood. When the blood has spread along 2/3 of the width, push the spreader slide forward with a steady even motion. The angle of the pusher slide and the speed of pushing may be changed to vary the thickness of the smear. The greater the angle, the thicker the smear. |
| - 5. | Label with computer-generated label. If label is not available, write the patient’s last name, first name and MRN or Accession and date on the frosted end of slide with a lead pencil. |
| - 6. | Allow the smear to air dry. |
| - 7. | Stain the smear using an automated slide stainer (Hematek). |

- III-Performing a Slide Review or Differential:

| Step | Action |
| --- | --- |
| 1 | Check slide identification. Check blood film identification to ensure that film and automated count report identification match. |
| 2. | Review automated results. Carefully review the automated report, noting the number of leukocytes and platelets, the number and size (MCV) of RBCs, the hemoglobin concentration (MCHC) of RBCs, and the RDW. During the slide review process, the information above is compared to the actual findings on the blood smear. Any discrepancy is investigated immediately and resolved before the differential is reported. |

| Step | Action |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 3. | Perform low power scan. Even Cellular Distribution WBC’s pulled to the feather end of the blood slide should not exceed 2-3 times the number of white cells present in the examination area of the smear. Remake slide which do not meet this standard. The examination area of the blood smear is typically the area where approximately 50% of the RBCs show minimal overlapping and 50% are individually spaced. Check Feather Edge for Fibrin Threads. Fibrin threads may be present even in films made from specimens that do not contain gross clots. Fibrin threads tend to entrap larger WBCs, and are accompanied by platelet clumping. Neither a differential nor a platelet estimate should be attempted on blood films that exhibit fibrin clots. Verify Stain Quality. The stain should clearly distinguish between dark purple nuclear material and bright red-orange erythrocytes. |  |  |  |  |  |  |  |
| - 4. | With the high power (40X) objective, observe several fields where RBCs are not overlapping. Cell distribution must be even. a) WBC Estimation - Perform a WBC estimation in the examination area of the blood film and compare to the value obtained by the Sysmex instrument. Any significant discrepancy must be investigated. - Accurate WBC estimates are dependent on having well made blood smears. Poor slide making technique can result in too many WBC’s being deposited on the feather or lateral edges. Reject smears having more than three (3) times the number of leukocytes on the edges than in the examination area. - With the high power (40X) objective, observe several fields where RBCs are not overlapping. - Count the number of white cells in six or more high power fields. - Determine the average number of white cells per high power field Average WBCs/hpf = Total number of WBCs Number of fields counted - From the table below, estimate the total white count: WBC estimation factors (use high power objective) |  |  |  |  |  |  |  |
|  |  |  | - # WBCs /HPF |  | - Estimated Total WBC Count |  |  |  |
|  |  |  | - 2-4 |  | - 4,000-7,000 |  |  |  |
|  |  |  | - 4-6 |  | - 7,000 – 10,000 |  |  |  |
|  |  |  | - 6-10 |  | - 10,000-13,000 |  |  |  |
|  |  |  | - 10-20 |  | - 13,000-18,000 |  |  |  |
| Step | Action |  |  |  |  |  |  |  |
| - 4. | b) Platelet Clumps or Other Abnormal Cells in the “Feather Edge” - The feather and lateral edges of the smear should be routinely examined in order to detect significant abnormalities not present in the smear area. - If platelet clumps are present on smear and count appears to be inaccurate: - Replace numerical platelet count with #DNR in WAM. #DNR will translate to “Clumping seen” in Cerner. - For numerical estimate go to step “5” - Note: The larger cellular elements tend to be pulled to the feather, leaving the - smaller WBCs (lymphocytes) in the body of the smear. Examine entire smear - for platelet clumps or evidence of fibrin, especially when validating an abnormally low platelet count. |  |  |  |  |  |  |  |
|  | - Rouleaux Formation and either WBC or RBC Agglutination - The presence of either rouleaux or agglutination should be indicated on the report or in comments section of WAM. |  |  |  |  |  |  |  |
| - 5. | With the 100X oil immersion objective, observe several fields where RBCs are not overlapping. |  |  |  |  |  |  |  |
|  | Platelet Estimation: |  |  |  |  |  |  |  |
|  |  | Platelets / oil field (avg of 10 fields) |  | Description |  | Estimated Platelet Count (thousands/mm3) |  |  |
|  |  | < 2.5 |  | Decreased |  | <50 |  |  |
|  |  | ≥ 2.5 but < 5 |  | Decreased |  | 50-99 |  |  |
|  |  | ≥ 5 but < 6.5 |  | Decreased |  | 100-129 |  |  |
|  |  | ≥ 6.5 but < 20 |  | Adequate |  | 130-400 |  |  |
|  |  | ≥ 20 but < 25 |  | Increased |  | 401-500 |  |  |
|  |  | ≥ 25 but < 30 |  | Increased |  | 501-600 |  |  |
|  |  | ≥ 30 |  | Increased |  | >600 |  |  |

| Step | Action |  |  |  |
| --- | --- | --- | --- | --- |
| - 5. | Platelet Morphology: |  |  |  |
|  | Examine for the following platelet morphological abnormalities using 100X OIF and grade as % of total platelets: FEW; MODERATE; MANY. Bizarre Platelets; 1 – 10%; 10 – 25%; >25%. Giant Platelets (>7 um); 1 – 10%; 10 – 25%; >25%. Large Platelets (4 - 6.5 um); 1 – 10%; 10 – 25%; >25%. Megakaryocytic; 1 – 2%; 3 – 5%; >5. Platelet Satellitosis; Grade Subjectively. Platelet Clumps; Grade Subjectively. |
|  | RBC Morphology Use guideline below to evaluate RBC morphology: Examine and grade RBC morphology using 100X OIF, approximating 200 RBC’s/field. Few; Moderate; Many. Acanthocytes; 1 – 2 cells; 3 – 5 cells; > 5 cells. Schistocytes; 1 – 2 cells; 3 – 5 cells; > 5 cells. Sickle Cells; 1 – 2 cells; 3 – 5 cells; > 5 cells. Spherocytes; 1 – 2 cells; 3 – 5 cells; > 5 cells. Tear Drop Cells; 1 – 2 cells; 3 – 5 cells; > 5 cells. Polychromasia; 1 – 2 cells; 3 – 5 cells; > 5 cells. Basophilic Stippling; 1 – 2 cells; 3 – 5 cells; > 5 cells. Siderocytes; 1 – 2 cells; 3 – 5 cells; > 5 cells. Pappenheimer Bodies; 1 – 2 cells; 3 – 5 cells; > 5 cells. Howell Jolley Bodies; 1 – 2 cells; 3 – 5 cells; > 5 cells. Helmet Cells; 1 – 2 cells; 3 – 5 cells; > 5 cells. Heinz Bodies; 1 – 2 cells; 3 – 5 cells; > 5 cells. Cabot Rings; 1 – 2 cells; 3 – 5 cells; > 5 cells. Hypochromasia; 10 – 50 cells; 51 – 100 cells; > 100 cells. Microcytosis; 10 – 50 cells; 51 – 100 cells; > 100 cells. Macrocytosis; 10 – 50 cells; 51 – 100 cells; > 100 cells. Ovalocytes; 10 – 50 cells; 51 – 100 cells; > 100 cells. Stomatocytes; 10 – 50 cells; 51 – 100 cells; > 100 cells. Target Cells; 10 – 50 cells; 51 – 100 cells; > 100 cells. Burr Cells; 10 – 50 cells; 51 – 100 cells; > 100 cells. Elliptocytes; 10 – 50 cells; 51 – 100 cells; > 100 cells. Crenated Cells; 10 – 50 cells; 51 – 100 cells; > 100 cells. Poikilocytosis*; 10 – 50 cells; 51 – 100 cells; > 100 cells. *Consider reporting actual RBC morphologies rather than Poikilocytosis.. |
|  | For Rouleaux and anisocytosis Few; Moderate; Many; Comments. Rouleaux; 1 – 25%; 25 – 50%; >50%; % of RBC forming rouleaux. Anisocytosis; 2X; 3 – 4X; >4X; Size of largest cells vs. smallest cell. |
| - 6. | Using either the HPF, 50 oil or 100 oil objective or Cellavision scan slide |  |  |  |
|  |  | - If: | - Then: |  |
|  |  | - No immature cells or abnormalities seen | - Validate the Automated Differential by Confiming the Cell Counter in CellaVision or by Clicking” Validate All” in WAM |  |
|  |  | ### Less than 10 Bands seen | - Validate the Automated Differential by Confiming the Cell Counter in CellaVision or by Clicking” Validate All” in WAM - . |  |
|  |  | ### Only 1 Metamyelocyte seen | - Validate the Automated Differential by Confiming the Cell Counter in CellaVision or by Clicking” Validate All” in WAM |  |
|  |  | - WBCs are abnormal or immature (younger than a band) | - Perform a 100 cell manual differential in WAM or by CellaVision. - For CellaVision, refer to LHS-2290 - Go to step 7 for manual differential in WAM. See also LHS-2320 |  |
| Step | Action |  |  |  |
| - 7. | Access WAM diff counter and perform manual differential. - Count and morphologically classify 100 leukocytes, in consecutive fields in strips or rows from the tail end straight back to the thick end of the smear. Each intact WBC must be classified. Skipping cells that look unfamiliar skews the results and is inaccurate and unacceptable practice. - If NRBCs enumerated during the manual differential do not match the instrument’s % NRBC counted, report instrument’s value. Instrument is more accurate than a manual differential count. - If 10 megakaryocytes /100 WBC are observed and WBC estimate from smear is lower than instrument result, perform a manual count and correct the WBC count using “Corrected WBC formula”. Corrected WBC = Uncorrected WBC X 100 100 + # of Megakaryocytes/100 WBC - Report any abnormal WBC morphology. - Report RBC morphology - Report Platelet Estimate |  |  |  |

White Blood        Examine for the following WBC morphological abnormalities and

Cell                       grade as “Few”, “Moderate”, or “Many” if present:

Morphology

- Atypical Lymphocytes

- Auer Rods

- Dohle Bodies

- Hypersegmentation

- Pelger-Huet

- Plasmacytes

- Reactive lymphocytes

- Sezary Cells

- Smudge Cells

- Toxic Granulation

- Vacuolated Granulocytes

- Vacuolated Nongranulocytes

| - Procedural Notes | - WBC abnormal scattergram: This IP message occurs whenever clustering in the WNR or WDF scattergrams is abnormal. Dashes may appear in place of data that was not calculated. If the analyzer has reported the WBC from the WDF channel, the WBC result will have the “&D” indicator adjacent to it. - If dashes ( - - ) appear in place of numeric data, reanalyze the sample, Scan the peripheral smear to verify the WBC, NRBC and differential results or perform a manual differential. - If asterisk ( * ) is next to the result, then Verify WBC, NRBC and differential result by scanning the slide for abnormal cells or platelet clumping and estimate the WBC and NRBC counts. Perform a manual differential if abnormal cells are observed. - If after reviewing the smear, and no abnormalities are found and the WBC and NRBC estimates match the analyzer reported WBC and NRBC, the results with astericks ( * ) may be reported. - If after reviewing the smear, and the NRBC count is significantly different than the analyzer count,( ie. XN = 0, Manual Diff = 30 ) and the WBC count is significantly different than the WBC smear estimate, perform a manual differential, enumerating the NRBC’s seen and correct the WBC count for the presence of NRBC’s counted using the formula: Corrected WBC = Uncorrected WBC X 100 100 + # of NRBC/100 WBC An uncorrected WBC count can be found in the Service/WNR tab of the Browser screen if needed. Use the TNC-N as the uncorrected WBC count. : |
| --- | --- |

- RETIC abnormal scattergram: This IP message indicates that the analyzer has detected increased activity in the RET-THR area or the RET-UPP area of the scattergram. RET-UPP is abnormal due to NRBCs, Howell-Jolly Bodies, parasites or stress reticulocytes. These are not included in the reticulocyte count. Asterisks ( * ) appear next to the RET%, RET#, IRF and RETHe parameters. These results may be unreliable and should be confirmed. - Prepare a 1:5 dilution with CELLPACK DCL diluent to minimize interference. Run the 1:5 dilution in the manual mode (NOT Predilute mode) Dilutions greater than 1:5 should not be used. NOTE: Do not use undiluted CELLPACK DST for any dilutions. - Check that the RBC (X5) on the diluted sample matches the original RBC count within 5% to ensure that the dilution errors have not occurred. Also, check that the diluted RBC count is not less than 0.05 x 10ⁿ/uL if result is less than 0.05 x 10ⁿ/uL then make a lower dilution to increase RBC count. - If the RET Abn Scattergram? Flag is eliminated, results may be reported. The Retic% and IRF do not need to be multiplied by the dilution factor since the percentages /ratios should remain the same regardless of the dilution factor. RETHe also does not need to be multiplied since it is measured at the cellular level and is unaffected by dilution. - If the flag is not eliminated or the RBC count is <0.50 x 10ⁿ/uL, review smear for the presence of polychromasia, parasites, NRBCs, Howell-Jolly bodies or basophilic stippling. If present, report the results with a comment saying that the results may be affected by the presence of interfering substances. :

|  | 2. Sources of error when making the slides are: - Dirty or greasy slides. - Excessive pressure of the spreader slide will cause margination of the leukocytes along the edges and at the end of the smear and may break fragile cells. - Touching the slide to the skin of the patient while obtaining the blood drop from the finger. - Excessive delay in spreading the blood may result in the blood clotting. |
| --- | --- |
|  | 3. Criteria for manually counting less than 100 cells: - WBC 1,500/mm3 or higher – count 100 cells - WBC less than 1,500/mm3 – count 50 cells Multiple smears (1-2) may be made in order to obtain a differential count of at least 50 cells. 50 cells should be counted to ensure that abnormal cells are not being missed. It is essential that the presence of malignant cells be detected on Oncology patients, as chemotherapy is often based on whether malignant cells are reported in the differential. Scan slides carefully for abnormal/immature cells. If observing two (2) peripheral smears fails to reveal a minimum of 50 WBCs, report differential in percentages of the number counted. 4. Refer to Sysmex® XN-Series Automated Hematology Systems Flagging Interpretation Guide for additional IP messages and interfering substances. |

| Author | Nanette Taft, CLS (ASCP) |
| --- | --- |

| ##### References | - Regional Hematology Scan, Differential, and Reporting Criteria Policy, June 23, 2004. - Maedel L, Sommers S, “Blood Cell Morphology”, ASCP Press, 1993. - Siene-Martin, E., “Clinical Hematology:Principles, Procedures, Correlations, 2nd Edition, 1998. - CAP competency assessment guidelines: RBC morphology grading system - Gulati GL:Blood Cell Morphology Grading Guide. Chicago, IL, ASCP Press, 2009 - Sysmex XN-Series Flagging Interpretation Guide Document number: 1166-LSS, Rev. 3 January 2017 |
| --- | --- |

Attachments           So.Cal Kaiser Permanente Regional WAM Rules- 03192018

| Related Documents | Documents | Title |
| --- | --- | --- |
|  | LHS 2520 | Sysmex DI-60™ |

Document History Page

Effective Date:  8/3/12

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Lab Manager reviewed/ date | Med. Director reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  | N.Taft 08/28/12 | N.Muneno 08/29/12 | S.Wirio, MD 08/30/12 | 08/30/12 |
| Major | Updated Platelet scan criteria to new Regional Critical Values and for Cerner. | N.Taft 07/23/14 | J.Wolf 07/24/14 | S.Wirio,MD 07/24/14 | 07/24/14 |
| Major | Updated for WAM rules changes 07/05/16; WBC026,030,034 | N.Taft 07/12/16 | J.Wolf 08/02/16 | S.Wirio,MD 08/02/16 | 08/02/16 |
| Major | Updated for WAM 5 regional rules |  |  |  |  |