| - Policy | - Each instrument and each piece of equipment in use in the hematology department will have daily, and/or weekly, and/or monthly maintenance performed as defined by manufacturer and/or Sherman Way Regional Laboratories Instrument Repair department guidelines and documented on designated preventive maintenance logs. - If any of the instruments or equipment fail it’s daily checks, it is removed from service, and Sherman Way Regional Instrumentation Repair, instrument manufacturer or engineering is notified. |
| --- | --- |

Preventive Maintenance Schedule Overview:

|  | - Instruments and Equipment | - Daily | - Weekly | - Monthly | - Other |
| --- | --- | --- | --- | --- | --- |
|  | XN-9000 | x | x |  |  |
|  | SP10 Stainer | x | x |  | As needed |
|  | CellaVision | x | x |  | As needed |
|  | ESR iSED |  |  |  | As needed |
|  | Eyewash |  | x |  |  |
|  | Safety Check | x |  |  |  |
|  | Diluent/Hemocytometer Check |  |  |  | As needed |
|  | Cytospin | x | x |  |  |
|  | Slide Stainer | x | x |  | As needed |
|  | CheckPoint Daily Reading Review | x |  |  |  |
|  | Shakers/Rotators | x |  |  |  |
|  | Manual Diff Counters | x |  |  |  |
|  | Multi-Purpose Timer | x |  |  |  |
|  | Microscopes | x |  |  |  |
|  | Vortex | x |  |  |  |

Preventive Maintenance Schedule Specifics:

XN-9000

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Shut-down Procedure - Start-up procedure | GY CLS |
|  | - Run XN-Check Controls | AM/PM/GY CLS |
| Weekly | - Power down entire XN Line | GY CLS |

SP-10 Stainer

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Shut-down Procedure ( Alternate 1 and 2 daily) - Start-up procedure - Clean Spreader Glass - Clean Stain Cassettes - Clean DI water container and refill with fresh DI H2O | GY CLS |
| As Needed | - Perform Rinse, Stain Line - Perform Rinse, Smear Line | GY CLS |

CellaVision

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - System Shutdown - System Startup - Clean X-Y Stage | GY CLS |
|  | - QC Localization Test (>97%) | AM CLS |
| Weekly | - Clean Objectives - Clear System Control Log | GY CLS |
| As Needed | - Replace Oil Immersion bag | AM/PM/GY CLS |

ESR iSED

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Run Seditrol QC | AM CLS |
| As Needed | - Replace printer paper (Green LED will flash) - Replace waste bottle ( Run wash cycle first ) - Replace iWASH bottle - Inspect interior surfaces for dust accumulation. Clean as needed - Checkrear fan for dust accumulation. Clean as needed. - Replace sample needle after 30,000 piercings (contact Tech Support). | AM/PM/GY CLS |

Eyewash

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Maintain clearance from obstruction | AM CLS |
| Weekly | - Flush water for 10 seconds or until clear. - Clear for adequate flow from both nozzles (flow should meet near top of arc in center). Note: Notify engineering if not working. | AM CLS |

Safety Checks

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - No Storage on the floor - No unlabeled reagents - Expiration and open date on all reagents. - Clean Bench Tops | AM CLS |

Hemacytometer and Diluent Checks

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| As Needed | - Check each new lot of Incyto Hemocytometers to ensure counting chambers are bright, clean and free of scratches. - Record lot # on action log. | AM CLS |
| As Needed | - Check each new lot # of sterile saline or KOH solution for particulate debris. - Record lot # on action log. | AM CLS |

Cytospin

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Clean spills and remove any debris. - Wipe down all surfaces. - Check for excessive noise or vibration. - Centrifuge must be balanced at all times. | AM CLS |
| Weekly | - Clean front panel with mild detergent. - Remove seal from rim of lid for cleaning and inspection; - Replace as necessary | AM CLS |

Slide Stainer:

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Clean platen and all external surfaces. - Empty waste tray. - Verify that drying fan is functional. - Visually check that all solutions are delivered to slide. - Verify proper operation by running slide. - Check slide for quality of stain. Note: If slides are chipped or broken during processing the instrument must be serviced. | AM CLS |
| Weekly | - Run methanol through stain cannula to platen | AM CLS |
| Monthly | - Replace pump and platen tubing as necessary | AM CLS |

Shaker/Rotators

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Wipe down all external surfaces - Verify proper operation. - Check for excessive noise or jerky motions. - Inspect state of rubber cup/sponge pad. | AM CLS |

Manual Diff Counter

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Wipe down surface. - Verify each key operation and alarm sounds at “100”. | AM CLS |

Multipurpose Timer

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Wipe down all surfaces. - Verify proper operation. | AM CLS |

Microscope

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily on all shifts | - Check condenser height adjustment. - Verify that field diaphragm is centered. - Verify that the condensor is centered . - Dry oil objective and condensor aftere each use of immersion oil. - Clean optics as required. - Check for any damage. - When not in use: - Turn lamp off (power). - Place dust cover over microscope. | AM/PM/GY CLS |

Vortex

| Frequency | Maintenance | Performed by: |
| --- | --- | --- |
| Daily | - Clean all external surfaces. - Verify proper operation. | AM CLS |

| Procedural Notes: | - Contact Sherman Way Regional Laboratories Instrument Repair: 8-397-7000 for troubleshooting instruments or equipment. - For the refrigerator and eyewash problems, contact the local engineering department: 8-324-4040 or eng beeper 11-1518 |
| --- | --- |

| Author | Mina Acosta, CLS, MT(ASCP) Nanette Taft, CLS, MT(ASCP) |
| --- | --- |

| Related Documents | Documents | Title |
| --- | --- | --- |
|  | LHS 2510 | Sysmex XN 9000 Procedure |
|  | LHS 1020 | Body Fluid Analysis |
|  | LHS 1062 | ESR: iSED |
|  | LHS 1110 | Wright Stain: Hematek 2000 Slide Stainer |

Document History Page

Effective Date:  08/16/04

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Lab Manager Reviewed/ Date | Med. Dir. Reviewed/ Date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| NEW | Change in instrument technology. | M.Acosta 10/04/04 | N.Muneno 10/07/04 | R.Doshi 10/12/04 | 10/12/04 |
| NEW | Change in instrument technology. | N.Taft 08/24/12 | N.Muneno 08/28/12 | S.Wirio, MD 08/30/12 | 08/30/12 |
| Major | Change in instrument technology. |  |  |  |  |