| Purpose | - This document will define the testing procedure, schedule and evaluation of quality control in the urinalysis department. |
| --- | --- |

| - Policy | - All unacceptable control results are resolved before testing patient specimens. - iChemVELOCITY Chemical QC is performed once on AM shift and once on PM shift when instrument is in use. - Advantus Chemical QC is performed once daily on AM Shift and whenever a new vial of 10SG Multistix is opened. - IRIS (iQ200) Microscopic QC is performed once on AM shift and once on PM shift when instrument is in use. - Binax Streptococcus Pneumoniae Ag QC is run once a day on the first sample of the day. |
| --- | --- |

| - Equipment - And - Reagents | ### Biorad Liquichek Urinalysis Control ### Multistix 10 SG Reagent Strips ### Advantus Urine Analyzer ### IRIS (iQ200/iChemVELOCITY) Urine Analyzers - iQ Control/Focus Set - IRIS Spec CA,CB,CC - Binax Streptococcus Pneumoniae Ag Test |
| --- | --- |

| Procedure |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | I. DAILY |  |  |  |  |  |  |
|  | AM Shift: |  |  |  |  |  |  |
|  |  | Test |  | Reagent | Instrument | Control |  |
|  |  | UR microscopic Exam (Manual) |  | N/A | Count-10 slides | Biorad Liquichek L1 & L2 |  |
|  |  | UR microscopic Exam (Automated) |  | Lamina | iQ200 | iQ Control/Focus Set |  |
|  |  | UR chemical Exam (Semi-automated Dipstick) |  | Multistix 10SG | Advantus | Biorad Liquichek L1 & L2 |  |
|  |  | UR chemical Exam (Automated) |  | iChemVELOCITY strips | iChem VELOCITY | IRIS Spec CA,CB,CC |  |
|  |  | UR Streptococcus Pneumoniae Test (SEE NOTE) |  | ### Binax Streptococcus Pneumoniae Kit | N/A | Binax Control Swabs Pos/Neg |  |
|  | - a. |  | - Bring controls to room temperature and gently mix before use. |  |  |  |  |
|  | - b. |  | - Record reagent lot #s and expiration dates in QC log. |  |  |  |  |
|  | - c. |  | - Perform tests in same manner as patient samples using the control material. |  |  |  |  |
|  | - d. |  | - Record results in designated area of QC log. |  |  |  |  |
|  | - e. |  | - Evaluate control results and resolve out of control runs before proceeding with patient testing. (see step # III of this procedure) - NOTE: Quality Controls (Binax Pos/ Neg Control Swabs) are run once a day on the first sample of the day and documented on the log sheets whenever you perform the test. If AM shift CLS did not get any patient sample on their shift then the PM/Night shift CLS needs to run QC on the first sample of the day. Samples are not tested unless controls are within acceptable limits. The Binax NOW Streptococcus pneumoniae Test has built-in procedural controls. Binax suggests that you record these controls, at minimum, for the first sample tested each day as well. |  |  |  |  |
|  | PM Shifts Daily QC: |  |  |  |  |  |  |
|  |  | Test |  | Reagent | Instrument | Control |  |
|  |  | UR microscopic Exam (Automated) |  | Lamina | iQ200 | iQ Control/Focus Set |  |
|  |  | UR chemical Exam (Automated) |  | iChemVELOCITY strips | iChem VELOCITY | IRIS Spec CA, CB, CC |  |
|  | - a. |  | - Bring controls to room temperature and gently mix before use. |  |  |  |  |
|  | - b. |  | - Perform test using the control material. |  |  |  |  |
|  | - c. |  | - Record results in designated area of QC log. |  |  |  |  |
|  | - d. |  | - Evaluate control results and resolve out of control runs before proceeding with patient testing. (see step # III of this procedure) |  |  |  |  |

|  | - II. UNSCHEDULED PARALLEL TESTING: |
| --- | --- |

|  |  | B. Multistix 10SG |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | - Multistix 10SG - Reagent strips |  |  |  |  |  | Biorad Liquichek - L1 & L2 |  |  |  |  |  |  |  |  |  |  | - Instrument |  |  |
|  |  |  | - Parallel testing new lot |  |  |  |  |  | √ |  |  |  |  |  |  |  |  |  |  | Advantus |  |  |
|  |  |  | - Each Shipment |  |  |  |  |  | √ |  |  |  |  |  |  |  |  |  |  | Advantus |  |  |
|  |  |  | - New Bottle opened |  |  |  |  |  | √ |  |  |  |  |  |  |  |  |  |  | Advantus |  |  |
|  |  | - Each shipment package received is labeled “DO NOT USE This Lot Number” and stored in the designated cabinet. - Indicate on form the lot number of Chemstrip being tested, lot number and expiration date of Biorad Liquichek controls, date of testing and initials of CLS tester. - A vial is tested using both levels of Biorad Liquichek controls. - If results are acceptable, place a label on each vial indicating that “This Lot is Ready for Use”. - If results are out, vial is retested. If repeat is still out, vial is discarded. Notation is made on form. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | - C. iChemVELOCITY STRIPS - iChemVELOCITY STRIPS; - IRIS SPEC CA, CB,CC; - Instrument. - Parallel testing new lot; √; iChemVELOCITY. - Each Shipment; √; iChemVELOCITY. |
|  |  | - Each shipment package received is labeled “DO NOT USE This Lot Number” and stored in the designated cabinet. - Indicate on form the lot number of iChemVELOCITY STRIPS being tested, lot number and expiration date of IRIS Spec CA, CB,CC controls, date of testing and initials of CLS tester. - A vial is tested using IRIS Spec CA, CB, CC controls. - If results are acceptable, place a label on each vial indicating that “This Lot is Ready for Use”. - If results are out, vial is retested. If repeat is still out, vial is discarded. Notation is made on form. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | E. iQ Control/Focus Set/ Lamina Reagent |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | iQ Control/Focus Set/Lamina; Instrument. - Parallel testing new lot; √; iQ200. - Each Shipment; √; iQ200. |
|  |  | - Each shipment package received is labeled “DO NOT USE This Lot Number” and stored in the designated cabinet. - Indicate on form the lot number and expiration date of Lamina or iQ Control/Focus Set controls, date of testing and initials of CLS tester. - A New iQ set or Lamina reagent is tested using current Lamina or iQ Control/Focus Set controls. - If results are acceptable, place a label on each vial indicating that “This Lot is Ready for Use”. - If results are out, set or reagent is retested. If repeat is still out, vial is discarded. Notation is made on form. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | F. Parallel Testing of New Reagent Lot #s |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 2. Biorad Liquichek L1 & L2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  | Control: Current Lot |  |  |  |  |  |  |  |  |  |  |  | Control: New Lot |  |
|  |  |  |  | Count-10 Method |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | A range for each method is established by performing 10-20 assays of new lot number. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3. Biorad Liquichek L1&L2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  | Control: Current Lot |  |  |  |  |  |  |  |  |  |  |  | Control: New Lot |  |
|  |  |  |  | Multistix 10SG |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | - Streptococcus Pneumoniae Ag Test Kit Rgt: Current Lot; Rgt: New Lot. Binax Positive Swab; ; . Binax Negative Swab; ; . - Binax Swab Controls Control: Current Lot; Control: New Lot. Streptococcus Pneumoniae Ag Test Kit; ; . |
|  | - III. Evaluating Control Run: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  | - Run Controls |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | - NO |  |  |  |  |  | - Within Acceptable Limits |  |  |  |  |  |  |  |  |  | - YES |  |  |  |  |  |
|  | - Document results in corrective action log. Circle and initial values. Check reagents for: outdate (age), correct lot number, and signs of contamination. Run controls again. |  |  |  |  |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. |  |  |  |  |  |  |  |  |
|  | - NO |  |  |  |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  |  | - YES |  |  |  |
|  | - Document results in corrective action log. Open a new vial of control. Run controls again |  |  |  |  |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. |  |  |  |  |  |  |  |  |
|  | - NO |  |  |  |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  |  | - YES |  |  |  |
|  | - Document results in corrective action log. Open a new container of reagent test strips and test with new control. |  |  |  |  |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. Discard bad vial of controls. |  |  |  |  |  |  |  |  |
|  | - NO |  |  |  |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  |  | - YES |  |  |  |
|  | - Document results in corrective action log. Troubleshoot instrument. Document on corrective action log. Run controls again. |  |  |  |  |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. Discard bad reagent test strips. |  |  |  |  |  |  |  |  |
|  | - NO |  |  |  |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  |  | - YES |  |  |  |
|  | - Document results in corrective action log. Get assistance from Regional Instrument Repair department. Notify supervisor. Use back-up instruments. Run controls on backup instruments if needed. Document results. |  |  |  |  |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. |  |  |  |  |  |  |  |  |
|  | - NO |  |  |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  |  |  | - YES |  |  |  |
|  | - Notify supervisor, and floors. Call instrument repair. |  |  |  |  |  |  |  |  |  |  |  |  | - Document findings. Proceed to testing patient samples. |  |  |  |  |  |  |  |  |

| - Procedural Notes | - All control testing is performed according to specific test procedure. |
| --- | --- |

| - Attachments | - Parallel Testing Log Sheets: - Multistix 10SG - iQ Lamina - Binax Strep. Pnuemo Test Kits - iQ Control/Focus Set - IRIS Spec CA,CB,CC |
| --- | --- |

| Author | Nanette Taft, CLS |
| --- | --- |

Document History Page

Effective Date:   12-21-02

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Lab Manager reviewed/ date | Med. Director reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  | M. Acosta 11-18-02 |  | R.Doshi,MD 12-21-02 | 12-21-02 |
| Revised | Incorporated safety message, discontinued SSA, Multistix 8SG, and Cards QS Preg. Added Genzyme. | C.D. Edora 10-09-06 | N. Muneno 10-09-06 | R.Doshi,MD 10-09-06 | 10-09-06 |
| Revised | Replaced Quantimetrix with Biorad Controls, discontinued weekly procedure and revised unscheduled procedure. | C.D. Edora 05-09-07 | N. Muneno 05-09-07 | R.Doshi,MD 05-09-07 | 05-09-07 |
| MAJOR | Updated procedure and replaced SUA with IRIS and added Strep pneumoniae. | C.D. Edora 08-18-08 | N. Muneno 08-18-08 | R.Doshi,MD 08-19-08 | 08-19-08 |
| MAJOR | Updated frequency of running the IRIS Spec CA & CB controls on AX4280 and Biorad L1 & L2 controls on CUA from 3X a day (every shift) to once a day. (AM shift only). Updated QC run for S. pneumoniae. Clinitest QC changed from daily to unscheduled testing. Discontinued Daily QC for Pregnancy (CLIA waived) kit. | Carlo Puno 03-17-09 | N.Muneno 03-18-09 | R.Doshi,MD 03-18-09 | 03-18-09 |
| MAJOR | Updated procedure and replaced CUA with Advantus, and Chemstrip 10 S-UA with Multistix 10SG. | Nanette Taft 09-30-10 | N.Muneno 10-04-10 | R.Doshi, MD 10-05-10 | 10-05-10 |

Document History Page

Continued

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Lab Manager reviewed/ date | Med. Director reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| MAJOR | Updated frequency of running CA,CB urine controls to be performed on each shift. | Nanette Taft 10/29/11 | N. Muneno 10/31/11 | R.Doshi MD 11/1/11 | 11-01-11 |
| MAJOR | Updated reagent lot verification procedrue for Lamina Reagent | N.Taft 04/16/12 | N.Muneno 04/21/12 | S.Wirio,MD 04/22/12 | 04-22-12 |
| MAJOR | Updated for New iChemVELOCITY Urine Analyzer. Changed frequency of Urine Controls to once daily for AM shift and once daily for PM shift. | N.Taft 01/15/15 | J. Wolf 01/16/15 | S.Wirio,MD 01/20/15 | 01-20-15 |
| MAJOR | Removed discontinued test and added to policy to run QC whenever a new vial of multistix is opened per insert |  |  |  |  |

See appropriate test procedure.