| ##### Safety Message | Use appropriate personnel protective equipment (PPE) while performing the studies. |
| --- | --- |

| ##### Policy | The Bacti department will perform correlation studies at least once every 6 months for all tests that have two or more methodologies and/or instruments available. |
| --- | --- |

| ##### Procedure |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Step | Action |  |  |  |  |
|  |  | 1. | A batch of at least 5 previously tested specimens or QC for both Rapid Flu and Rapid RSV are run on all instruments: Note: Specimens must not have exceeded the required acceptable stability limits for tests. |  |  |  |  |
|  |  |  |  | Analyzer or Method #1 | Analyzer or Method #2 |  |  |
|  |  |  |  | Veritor #1 for FLU | Veritor #2 for FLU |  |  |
|  |  |  |  | Veritor #1 for RSV | Veritor #2 for RSV |  |  |
|  |  | 2. | A log of results is maintained. A passing criterion of 100% of total values correlating with the correlation counterpart must be met. |  |  |  |  |
|  |  | 4. | Laboratory Director or designee will review correlation for clinical significance. |  |  |  |  |
|  |  | 5. | For tests that fail correlation study, a review of instrument quality control, maintenance, and calibration records is performed. Regional Instrument Repair department or instrument manufacturer may be consulted. |  |  |  |  |

| ##### Author | Nanette Taft, CLS, MT (ASCP) |
| --- | --- |

| ##### Reference | - Accreditation Manual for Pathology and Clinical Laboratory Services, Standard PD.5.3.3.1 1996 CAMPCLS, Quality Control, Q.C.1.2.1. |
| --- | --- |

Document History Page

Effective Date: _04/24/2018_____________

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Lab Operations Director Reviewed/ Date | Medical Director Reviewed/ Date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  |  |  |  |  |