| # PURPOSE | This procedural document provides instructions on how to initiate and suspend autovalidation and autoverification in Sysmex XN 9000. |
| --- | --- |

| # POLICY | Autovalidation and autoverification should be suspended anytime that a component of the system or any of the algorithms is found to be out of order. |
| --- | --- |

| # SCOPE | All laboratory personnel who work on systems with autovalidation/autoverification capabilities have the access and right for this document usage whereby circumstances need them to for job duties. |
| --- | --- |

| # SAFETY PRECAUTIONS | Refer to the safety manual for general safety requirements. |
| --- | --- |

PROCEDURE:

|  | # Step | # Action |  |
| --- | --- | --- | --- |
| # Suspension of Autovalidation | # 1. | Suspend autovalidation if: - QC results are unacceptable even after thorough investigation and corrective action - Instrument malfunction - Failed calibration - Reagent problem - Any other situation where instrument results are suspected to be adversely affected, e.g. high temperature in the lab - Consult lead CLS/Supervisor/Sysmex TAC if unsure |  |
|  | # 2. | To suspend autovalidation on XN-9000 - Follow regional procedure for Turning Autoverification On and Off in Cerner PathNet located in LabNet - Use service resource HAR SysWAM5 MP ss |  |
|  | # 3. | Document on daily checklist. |  |
|  | # Step |  | # Action |
| # Re-initiation of Autovalidation | # 1. |  | Re-initiate autovalidation if: - QC results are acceptable - Instrument problem has been corrected - Calibration has passed - Reagent problem has been solved - Any other situation has been corrected and instrument results are now acceptable - Consult lead CLS/Supervisor/Sysmex TAC if autovalidation/autoverification can be re-initiated |
|  | # 2. |  | To re-initiate autovalidation on XN-9000 - Follow regional procedure for Turning Autoverification On and Off in Cerner PathNet located in LabNet - Use service resource HAR SysWAM5 MP ss |
|  | # 3. |  | Document on the daily checklist. |

# REFERENCES:

Turning Autoverification On and Off in Cerner PathNet:

Document No. SCPMG LIS-0083

# AUTHOR:

# Nanette Taft, CLS(ASCP)

HISTORY PAGE

Effective Date: 08/28/18

| Type of Change: New, Major, Minor | Description of Change(s) | Area Lab Manager/Date | Laboratory Director Review/Date | Medical Director Review/Date | Date Change Implemented |
| --- | --- | --- | --- | --- | --- |
| New |  |  |  |  |  |