Limited Urinalysis

Roche Chemstrip-9

The CHEMSTRIP urine testing system is a multi-parameter test strip to measure certain constituents in the urine.  These measurements are useful in the evaluation of renal, urinary, and metabolic disorders.

SUMMARY:

CHEMSTRIP urine test strips are inert plastic strips to which are attached different reagent papers for indication of leukocytes and determining nitrite, pH, protein, glucose, ketones, urobilinogen, bilirubin, and blood in urine.  Refer to the outside box and vial label for the specific tests.  The test papers are uniquely attached to the strip with a nylon mesh which holds the reagent paper in place, protects the paper, and provides for rapid and even wetting of the entire test area.  To prevent urine runover, certain of the test papers have an inert absorbent paper located between the test area and the strip.

The CHEMSTRIP urine test strips are packaged in an aluminum vial with a tightly fitting cap which contains a drying agent.  Each test strip is stable and ready for use when removed from the vial.  No additional instrumentation is required.

TEST PRINCIPLE--See Package Insert

REAGENTS:

Leukocytes--

Indoxylcarbonic acid ester                 5.8  micro gm

Diazonium salt                             2.1  micro gm

Buffer                                   725.0  micro gm

Nonreactive ingredients                  641.7  micro gm

Nitrite--

3-Hydroxy-1,2,3,4-tetrahydro-7,8

benzoquinoline                          12.50 micro gm

Sulfanilamide                             11.00 micro gm

Buffer                                   157.50 micro gm

pH--

Bromthymol blue                            5.25 micro gm

Methyl red                                 0.42 micro gm

Phenolphthalein                            2.63 micro gm

Nonreactive ingredients                  431.67 micro gm

Protein--

3',3",5',5"-Tetrachlorophenol-3,4,5,6-

tetrabromosulfophthalein                 4.88 micro gm

Buffer                                   829.17 micro gm

Nonreactive ingredients                  201.67 micro gm

Glucose--

Tetramethylbenzidine                      38.3  micro gm

Glucose oxidase                            1.92 U

Peroxidase                                10.33 U

Buffer                                    74.17 micro gm

Nonreactive ingredients                 1481.7  micro gm

Ketones--

Sodium nitroferricyanide                  55.33 micro gm

Glycine                                    1.60 micro gm

Buffer                                     3.29 micro gm

Nonreactive ingredients                  201.67 micro gm

Urobilinogen--

4-Methoxybenzene-diazonium-

tetrafluoroborate                       25.00 micro gm

Buffer                                   981.67 micro gm

Nonreactive ingredients                   35.00 micro gm

Bilirubin--

2,6-Dichlorobenzene-diazonium-

tetrafluoroborate                        7.50 micro gm

Buffer                                  1215.83 micro gm

Nonreactive ingredients                   31.67 micro gm

Blood--

Tetramethylbenzidine                      20.83 micro gm

2,5-Dimethyl-2,5-dihydroperoxyhexane     133.33 micro gm

Buffer                                   911.67 micro gm

Nonreactive ingredients                 1460.00 micro gm

Warnings and precautions:

For In Vitro Diagnostic Use.

Warning. Toxic.  CHEMSTRIP urine test strips contain one or more of the following chemicals: phenol, diazonium salt, nitroferricyanide.  Avoid contact with skin and mucous membranes; flush affected areas with copious amounts of water.  Get immediate medical attention for eyes or if ingested.  Exercise the normal precautions required for handling all laboratory reagents.

Storage and Stability:  Store at temperatures under 300C (860).  Do not freeze.  CHEMSTRIP urine test strips are stable in the original capped vial until the listed expiration date.  In order to avoid exposure to moisture, the vial must be closed immediately after removal of a strip, using the original stopper which contains a drying agent.

Specimen Collection and Preparation:  CHEMSTRIP urine test strips may be used on any freshly voided urine specimen or on urines collected under special conditions, such as first-morning specimens and post-prandial urines.  The urine must be collected in a clean container and should be tested as soon as possible after collection (do not centrifuge or use preservatives).  It is of particular importance to use fresh urine to obtain the best results with the test for urine bilirubin, as this compound is very unstable when exposed to room temperature and daylight.  If testing cannot be performed within one hour after collection, the specimen should be refrigerated at 2 to 80C immediately and returned to room temperature before testing.  Mix thoroughly before use.  Urine should be collected in a container which will allow complete immersion of the reagent areas on the test strip.  If a cleanly voided urine is not collected, a positive test result for leukocytes may be due to a source of leukocytes external to the renal-urinary system.

PROCEDURE:

Materials Provided:  1 aluminum vial containing 100 CHEMSTRIP urine test strips.  A visual comparison color scale for reading test results is printed on the vial label.

Materials Required:  Biorad urinalysis for quality control check, timing device.

Assay:

1.  Briefly (no longer than 1 second) dip test strip into the urine.  Ensure that the chemically impregnated patches on the test strip are totally immersed.

2.  Draw the edge of the strip along rim of specimen container to remove excess urine.

3.  After the appropriate time, read the test as follows:

Hold strip close to color blocks and match carefully, ensuring that the strip is properly oriented to the color chart on the vial label.

Leukocytes...................................60-120 seconds

Nitrite..........................................30 seconds

pH................................................immediate

Protein.......................................30-60 seconds

Glucose..........................................60 seconds

Ketones..........................................60 seconds

Urobilinogen..................................10-30 seconds

Bilirubin.....................................30-60 seconds

Blood........ ...................................60 seconds

For convenience, all values on the strip may be read between 1 and 2 minutes after immersion in the urine.  The colors are stable for up to 120 seconds after immersion.  Color changes that occur after 2 minutes from immersion are not of diagnostic value.  Color changes that occur only along the edge of the test area should be ignored.  Careful removal of excess urine (step 2) should eliminate this effect.

Calibration:

Calibration of the CHEMSTRIP 9 urine test strips by the user is not required.

Quality Control:

BioRad Liquichek: Liquid solutions of human urine are tested as unknown samples, level 1 and 2,  with assayed values defined for each lot on a value sheet for physical and chemical constituents. When Liquichek is stored at 2-8 C, the constituents are stable for a maximum of 30 days after opening.

Two levels of control are run each day of testing, and whenever a new vial of strips are opened.

Control results must be within the stated ranges on the value sheet.  If discrepancies arise:

1.  Ensure that the reagent strips have not become discolored by exposure to air.

2.  Ensure good saturation of the pads with the control, then blot the strip on a paper towel to prevent run off /bleeding of reagents from pad to pad.

3.  If values remain beyond the stated range, try a different container of stips, then a different lot number of strips.

4.  If repeat testing fails to resolve discrepancies, call the Ancillary Testing Coordinator  at 541 826-2111 X 3251.

RESULTS:

Results are obtained by direct visual comparison with the color scale printed on the vial label.  No calculations are necessary.

The leukocyte visual color chart bears four color patches.  The first patch, designated as negative, indicates the absence of leukocytes.  Negative results should be read at 60 seconds and verified at 120 seconds.  The second patch, designated as trace, indicates a possible borderline situation, and it is recommended that the test be repeated on a fresh urine sample from the same patient.  The next two patches, designated as + and ++, indicate the presence of increasing amounts of clinically significant numbers of leukocytes.  Positive and repeated trace findings indicate the need for further testing of the patient and/or urine sample in accordance with the medically accepted procedures for pyuria.

The leukocyte visual color chart is not intended to represent quantitative findings and serves only as a screening mechanism for the presence of leukocytes in urine.  If quantitative results are desired, it is recommended that microscopic examination of the urine be carried out utilizing a reference procedure.

LIMITATION:

The limitations including interfering substances for each reagent are shown below.

Leukocyte Test:  This test is not affected by erythrocytes in concentrations up to 10,000/microL or by bacterial common in urine.

Specimens should not be collected in containers that have been cleaned with strong oxidizing agents. Do not use preservatives.

The drugs cephalexin and gentamicin have been found to interfere with this test.  In addition, nitrofurantoin colors the urine and this effect interferes with visual interpretation of the test strip.  High levels of albumin (> 500 mg.dl) in the urine may interfere with the test results.

Nitrite Test:  Large amounts of ascorbic acid (see under glucose) decrease the sensitivity of the test.  False-positive readings may be produced by medication that colors the urine red or which turns red in an acid medium (e.g., phenazopyridine).

pH Test:  No known interferences when handled according to instructions.

Protein Test:  False-positive results may be found:

1.  In strongly basic urine (pH 9 or higher).

2.  During therapy with phenazopyridine.

3.  When infusions of polyvinylpyrrolidone (blood substitutes are administered.

4.  When residues of disinfectants containing quaternary ammonium groups or chlorohexidine are present in the urine container.

Glucose Test:  The effect of ascorbic acid (vitamin C) retained in the urine due to ingestion of vitamin tablets, antibiotics or fruit juices has been eliminated at glucose concentrations of 100 mg/dl and above so that false-negative readings may only rarely occur, even at high concentrations of ascorbic acid.  False-positive readings may be produced by strongly oxidizing cleaning agents in the urine container.

Ketone Test:  Red-orange to red color shades, which are, however, readily distinguished from the colors obtained with ketone bodies, can be produced by phenylketone or phthalein compounds that may be administered for liver and kidney function tests.

2-Mercaptoethane sulphonate sodium (MESNA) or other sulfhydryl-containing compounds may cause false-positive results

Urobilinogen Test:  The total absence of urobilinogen cannot be detected.  Most normal urines give a light pink reaction.  The test gives the same color reaction with urobilinogen as with stercobilinogen; however, the differentiation is not of diagnostic importance.  Urine from patients who are being treated with phenazo1pyridine may show a false-positive reaction.  Nitrite concentrations above 5 mg/dL or formalin concentrations above 200 mg/dL (as a preservative) may cause a decrease in the color reaction.

Bilirubin Test:  Large amounts of ascorbic acid present in urine following the ingestion of medication containing vitamin C or fruit juices lower the sensitivity of the test.  In case of doubt, the test should be repeated on urine voided at least 10 hours after the last administration of vitamin C.  Elevated concentrations of nitrite, as in urinary-tract infections, may result in lower bilirubin values.  Large amounts of urobilinogen in the urine affect the color change of the bilirubin test, but not enough to give a positive result.  False-positive readings may be produced by a medication that colors the urine red, or which turns red in an acid medium (e.g. phenazopyridine).

Blood Test:  False negative readings are obtained when formalin is used to preserve the urine.  Nitrite in excess of 10/mgdL in the urine (which is rare in urinary-tract infections) delays in the reaction.  False-positive results can be produced by residues of strongly oxidizing cleaning agents in the urine container.  Urine from menstruating females will occasionally yield a positive result.  This test has not been found to be affected by the ingestion of reasonable quantities of ascorbic acid.

EXPECTED VALUES:

Leukocytes:  Normal urines should produce no color reaction.  Positive and repeated trace findings indicate the need for further testing of the patient and/or urine sample in accordance with the medically accepted procedures for pyuria.

Nitrite:  A concentration as low as .05 mg/dL of nitrite will produce a slightly pink coloration of the test area.  This indicates a positive result.

pH:  Urine pH values generally range from 5 to 9 units.  The most frequent pH values for the first morning specimens in healthy subjects are between pH 5 and 6.

Protein:  A color change from yellow to light green/green will occur if protein is present in urine.  The concentrations given on the vial label correspond favorably with the albumin concentration in urine.  Pathological proteinuria will usually produce persistent values above 30 mg/dL.  Clinical significance of the trace result should be determined by additional testing.

Glucose:  Due to the test's sensitivity, glucose should not be detectable in normal urine.  Therefore, any positive reaction should be followed by further diagnostic evaluation of the patient, such as a quantitative blood glucose or a glucose tolerance test.

Ketones:  Ketone bodies should not be detected in normal urine with this test.  Fasting or starvation diets may cause positive indications.  In known pathological conditions such as diabetes, the presence of ketones may be useful as an index of metabolic status.

Urobilinogen:  Concentrations are usually greater in the afternoon than during the remainder of the day.  Values up to 1 mg/dL are usually considered normal.

Bilirubin:  In normal urine, bilirubin should not be detectable with this test.  However, the test is very sensitive to bilirubin (0.5 mg/dL will produce positive results) and any positive reaction indicates that further diagnostic evaluation of the patient is needed.

Blood:  Erythrocyte excretion up to 5 Ery/microL may be expected in normal urine.  Levels above these certainly warrant further diagnostic evaluation of the patient.

REPORTING & CONFIRMING

Results are to be entered into the patients CPRS Progress Notes using the title Lab Note using the urine dipstick template. All specimens screened at the K-Falls CBOC are to be poured off into a Boricon transfer container within 2 hours of collection, positively identified with an appropriate Order Label and collection time, and shipped to the White City laboratory for complete urinalysis.