VA SORCC

CBOC OCCULT BLOOD TESTING

FECAL OCCULT BLOOD (Beckman-Coulter Hemoccult SENSA):

PURPOSE:

The Hemoccult SENSA test is a rapid qualitative unmodified guaiac slide test for the qualitative detection of fecal occult blood which may be indicative of gastrointestinal disease.  It is not a test of colorectal cancer or any other specific disease.  The Hemoccult SENSA test is recommended for professional use as a diagnostic aid during routine physical examinations; to monitor for gastrointestinal bleeding in patients with iron deficiency anemia, peptic ulcer, ulcerative colitis and other conditions, or who are recuperating from surgery; and in screening programs for colorectal cancer when the Patient Instructions are closely followed.  The Hemoccult SENSA test and other unmodified guaiac tests are not recommended for use with gastric specimens.

The detection of occult blood is critical to many gastrointestinal diseases.  The presence of occult blood in fecal material may indicate gastrointestinal pathology such as hemorrhoids, diverticulitis, fissures, colitis or colorectal cancer.  Fortunately, these conditions can be detected with several diagnostic methodolgies available which include testing of stools for occult blood, complete physical examination with digital examination, and proctosigmoidoscopy.  Air contrast barium enema and fiberoptic colonoscopy also contribute significantly to the diagnosis of colonic lesions.  Unfortunately, only a small percentage of bowel and rectal cancers are found on digital examination and patients with no symptoms of bowel disease do not readily present themselves for procedures such as proctosigmoidoscopy and barium enema.

When used for colorectal cancer screen, the American Cancer Society recommends testing of at least three consecutive bowel movements due to the irregularity of bleeding from lesions in the intestinal tract.  Serial fecal specimen analysis is also recommended when screening asymptomatic patients.

If a positive result is obtained with the test, a follow-up with additional diagnostic tests, as soon as possible, is essential.  With any occult blood test, results cannot be considered conclusive evidence of the presence of or absence of gastrointestinal bleeding or pathology.  The test is not intended as a replacement for other diagnostic procedures such as proctosigmoidoscopy examination, barium enema, and X-ray studies.

PRINCIPLE:

Van Deen is generally credited with the discovery that gum guaiac, a natural resin extracted from the wood of Guiacum officinale, is useful in detecting occult blood.

The Hemoccult SENSA test more reliably detects abnormal bleeding associated with gastrointestinal disorders than standard guaiac tests.  As a result, it will have a higher sensitivity for disease but also a higher false positive rate among non-diet compliant patients.  Hemoccult SENSA positive test results appear as more stable, intense blue color reactions than the results of other guaiac tests, improving overall readability and precision.  As with other guaiac tests, accuracy depends upon the status of the patient at the time the specimen is taken and may be affected by interfering substances.

The Hemoccult SENSA test is a simplified and standardized variation of the laboratory guaiac procedure for detection of occult blood. The Hemoccult SENSA test formulation includes an enhancer which makes the test more sensitive and more readable than other guaiac-based tests.  Because the Hemoccult SENSA test requires only a small fecal specimen, offensive odors are minimized and storage or transport of large fecal specimens is unnecessary.

The Hemoccult SENSA test is based on the oxidation of guaiac by hydrogen peroxide to a blue-colored compound.  The heme portion of hemoglobin, if present in the fecal specimen, has peroxidase activity which catalyzes the oxidation of alpha guaiaconic acid (active component of the guaiac paper) by hydrogen peroxide (active component of the developer) to form a highly conjugated blue quinone compound.

## REAGENTS AND MATERIALS:

The Hemoccult SENSA test consists of two main components:

  Hemoccult SENSA Slides (Test Cards) containing guaiac paper

  Hemoccult SENSA Developer, a developing solution containing a stabilized mixture of

less than 4.2% hydrogen peroxide, 80% denatured ethyl alcohol and enhancer in an aqueous

solution.

Hemoccult SENSA Single Slides are convenient for use when single fecal specimens are tested.

Hemoccult SENSA Slide (blue striped, product #64151)

Hemoccult SENSA Developer (blue cap and striped label, prodsuct #64115)

Applicator sticks.

Precautions:

  For in vitro diagnostic use.

  Do not use after expiration date which appears on each test component.

  Because this test is visually read and requires color differentiation, it should not be

interpreted by individuals with blue color deficiency (blindness).

  Patient specimens, and all materials that come in contact with them, should be handled as

potentially infectious and disposed of using precautions.

  Slides (blue and green striped)

Keep cover flap of slide sealed until ready to use.  Protect slides from heat, light, and volatile

chemicals. Hemoccult SENSA  slides present no hazard to the user.

  Developer (blue and green striped label with blue bottle cap)

Hemoccult SENSA Developer should be protected from heat and the bottle kept tightly

capped when not in use.  It is flammable and subject to evaporation.

·  Hemoccult SENSA Developer is an irritant.  DO NOT USE IN EYES.  AVOID CONTACT

WITH SKIN.  Should contact occur, rinse promptly with water and consult a physician.

IMPORTANT:  Use Developer (blue and green striped label with blue bottle cap) only with Hemoccult SENSA Slides.  Do not interchange Hemoccult SENSA with Hemoccult test reagents, which are identified by yellow and green striped packaging, or with components from any other manufacturer.

Storage & Stability:

Store product at controlled room temperature (15 to 30C) in original packaging.  Do not refrigerate or freeze.  Protect from heat and light.  Do not store with volatile chemicals (e.g., iodine, chlorine, bromine, or ammonia).  The Hemoccult SENSA Slides and Developer will remain stable until the expiration dates which appear on each side and developer bottle when stored as recommended.

SPECIMEN COLLECTION AND HANDLING:

The hemoccult SENSA test requires only a small fecal specimen.  The specimen is applied to the guaiac paper by the attending physician.  As immediate testing is required, wait 3 to 5 minutes to allow the specimen to penetrate the paper before developing.  Open the back of the slide and apply two drops of Hemoccult SENSA Developer to guaiac paper directly over each smear.  Read results within 60 seconds.

Any trace of blue on or at the edge of the smear is positive for occult blood.

Developing the Quality Control Feature:

The Performance Monitor area must be developed on every slide.  Apply one drop of Hemoccult SENSA Developer between the positive and negative areas.

Read results within 10 seconds.  If the slide and developer are functional, a blue color will appear in the positive area and no blue will appear in the negative area.

Neither the intensity nor the shade of the blue from the positive Performance Monitor area should be used as a reference for the appearance of positive test results.

Any blue originating from the positive Performance Feature should be ignored when reading the sample test results.

If there is no color development with the positive are, or a blue color development with the negative area, contact the lab.

RECORDING THE RESULTS:

Enter the test results into the text of the patient’s progress note as soon as possible after testing.  Results in Klamath Falls will be entered into Vista Lab using the “White City Lab ByPass” option.

# INTERFERING SUBSTANCES:

Refer to patient instructions in “Specimen Collection & Handling” section.  In general, patients should be carefully instructed to not ingest foods and vitamins which can cause false-positive or false-negative test results for at least 72 hours before and through the test period.

Substances which can cause false-positive test results:

  Red meat (beef, lamb and liver)

  Aspirin (greater than 325 mg/day) and other non-steroidal anti-inflammatory drugs such

as ibuprofen, indomethacin and naproxen.

  Corticosteroids, phenylbutazone, reserpine, anticoagulants, antimetabolites, and cancer

chemotherapeutic drugs.

  Alcohol in excess.

  The application of antiseptic preparations containing iodine (povidone/iodine mixture).

Dietary iron will not produce false-postive test results with Hemoccult SENSA tests; however, vitamin C included with some iron supplements may cause a false positive result.

Acetaminophen  is not expected to affect test results.

Substances which can cause false-negative test results:

  Ascorbic acid (vitamin C) in excess of 250 mg per day.

  Excessive amounts of vitamin C enriched foods citrus fruits and juices.

  Iron supplements which contain quantities of vitamin C in excess of 250 mg per day.

Discolorations:

Occasionally, a light blue discoloration may be noticed on the guaiac test paper.  This discoloration does not affect the accuracy or performance of the test when it is developed and interpreted according to the recommended procedure.  When developer is added directly over the fecal smear on a discolored slide, the blue background color migrates outward.  A blue ring forms at the edge of the wetted area, leaving the guaiac paper around the fecal smear off-white in color.  Any blue on or at the edge of the smear is postive for occult blood.  Proper storage of Hemoccult SENSA Slides will help prevent blue discoloration.

Some specimens have a high bile content which causes the feces to appear green.  A distinct green color (no blue), appearing on or at the edge of the smear within 60 seconds after adding Hemoccult SENSA Developer should be interpreted as negative for occult blood.  A blue or blue-green color should be interpreted as positive for occult blood.

PERFORMANCE MONITOR FEATURE (QUALITY CONTROL):

The function and stability of the slides and developer can be tested using the on-slide Performance Monitor feature.  The positive (+) and negative (-) Performance Monitor areas are located under the sample area on the developing side of the slides.

The postive Performance Monitor area contains a hemoglobin-derived catalyst which will turn blue within 10 seconds after applying developer.  The negative Performance Monitor area contains no such catalyst and should not turn blue after applying developer.

The Performance Monitor feature provides assurance that the guaiac paper and developer are functional.  In the unlikely event that the Performance Monitor areas do not react as expected after applying developer, the test results should be regarded as invalid.  Should this occur, contact the Ancillary Testing Coordinator at White City for assistance (X3632).

Hemoccult SENSA Monitors must be tested, and results recorded, with every new lot of Hemoccult SENSA Slides and every new lot of Hemoccult SENSA Developer.

LIMITATIONS OF PROCEDURE:

Bowel lesions, including some polyps and colorectal cancers, may not bleed at all or may bleed intermittently.  Also, blood, if present, may not be distributed uniformly in the fecal specimen.  Consequently, a test result may be negative even when disease is present.

Conversely, a Hemoccult SENSA test result may be postitive on specimens from healthy patients.  This may be due to interfering substances in the diet or to medications.  It may also be due to low but detectable levels of blood loss, common to both healthy adults and patients with gastrointestinal disease.

Therefore, as with any occult blood test, results with the Hemoccult SENSA test cannot be considered conclusive evidence of the presence or absence of gastrointestinal bleeding or pathology.  Hemoccult SENSA tests are designed for preliminary screening as an aid to diagnosis.  They are not intended to replace other diagnostic procedures such as sigmoidoscopy, colonoscopy, barium enema, or other x-ray studies.

The Hemoccult SENSA test, as well as other unmodifed fecal occult blood tests, should not be used to test gastric specimens.  Interfering factors, such as low pH, high drug concentrations, metal ions or plant peroxidase in food, may affect the function of guaiac-based occult blood tests.  Gastroccult, available from Beckman Coulter Primary Care Diagnostics, is specifically designed to detect occult blood in gastric specimens.

Addition of a drop of water (rehydration) to the guaiac test card prior to the addition of the developer increases the sensitivity of the test, but also increases the number of false-positive test results.  For this reason, rehydration is not a recommended procedure for the Hemoccult SENSA test.

EXPECTED RESULTS:

In a general screening population of highly compliant asymptomatic individuals, a positivity rate of approximately 3% was obtained; among a similar group of less compliant individuals, a positivity rate of about 7% was observed.  The false-positivity rate for colorectal disease was   1 to 3% depending on the population studied.

Positivity rates for fecal occult blood tests have been shown to vary in each patient population depending on diet, age, predisposition to colorectal disease, and other factors that may be associated with bleeding gastrointestinal lesions.

## NORMAL VALUES:

Negative.

PERFORMANCE CHARACTERISTICS:

Early detection of colorectal cancer in asymptomatic, average risk individuals is neccesary to reduce mortality.  The detection of occult blood in stool using Hemoccult SENSA, a more sensitive and readable test than Hemoccult, is a highly effective method for detecting bleeding associated with colorectal cancer.

Clinical studies using [51Cr] chromium – labeled blood cells suggest that a daily blood loss of  2-3ml (appoximately 0.3 mg hemoglobin/gm feces) is the lower limit of blood loss that may be associated with gastrointestinal pathology.  Based on in vitro studies in which fecal samples from asymptomatic, normal volunteers were spiked with fresh whole blood, Hemoccult SENSA gave positive test results about 75% of the time at 0.3 mg Hb/gm feces.  The positivity rates increased as the equivalent daily blood loss increased.  Virtually all Hemoccult SENSA tests were positive at an equivalent daily blood loss equal to or greater than 10mL.

The specificities of the Hemoccult SENSA test and the Hemoccult test are the same when normal subjects follow the Patient Instructions as recommended.  When normal subjects consume large amounts of red meat, the false-positive rate will be higher with the Hemoccult SENSA test.  Raw fruits and vegetables in the diet give about the same number of false-positive test results when samples are tested immediately after collection.  To reduce the rate of false-positive tests due to consumption of raw fruits and vegetables, the Hemoccult SENSA test is best developed three days after sample application.  This three day delay allows for degradation of any fruit and vegetable peroxidases that may be present in the fecal sample.

For more information regarding the clinical performance of the Hemoccult SENSA test, please refer to the package insert.

REFERENCES:

Hemoccult SENSA product insert; revised October 2002, Beckman Coulter, Inc.

Prepared By:

| Adopted |  |  |
| --- | --- | --- |
| Reviewed |  |  |
| Reviewed |  |  |
| Reviewed |  |  |
| Reviewed |  |  |
| Reviewed |  |  |
| Reviewed |  |  |
| Reviewed |  |  |
| Reviewed |  |  |
| Reviewed |  |  |
| Revised |  |  |

Supercedes: