Accu-Chek Inform II CBG Procedure

I.	PRINCIPLE/PURPOSE

Any patient requiring blood glucose levels for the purpose of monitoring and managing diabetes control or glucose levels. A review of glucose metabolism is essential in order to appreciate blood glucose concentration variations that may reflect primary abnormalities of carbohydrate metabolism as well as secondary abnormalities accompanying other diseases.

II.	MATERIALS

A. Gloves

B. Accu-Chek Safety Pro Lancets (One time use devices)

C. Alcohol pads

D. Gauze or cotton balls

E. Accu-Chek Inform II test strips (Stored in Lab)

D. Accu-Chek Inform II Control Levels 1 & 2 (Stored in Lab)

III.	EQUIPMENT

A. Accu-Chek Inform II CBG Meter

B. Computer with access to CPRS/Vista

IV.	QUALITY CONTROL

A.PURPOSE

Quality control testing validates the integrity of the strips, the correct coding and 	     	    	calibration of the meter, and operator technique. Therefore, it should be done on a 		routine basis and whenever there is a change in the meter, strips or with 			      	questionable test results. The meter will alert the user when the QC testing is due. 	  	      	If QC is not performed, the meter will lock-out further patient testing until QC is 		      	performed and an acceptable result is obtained.

B. MATERIAL STABILITY

1. Accu-Chek Inform II control levels 1 and 2 are stable for 90 days once 				opened at temperatures between 4 to 30°C. A new expiration date must be 				written on the bottles once they have been opened. If the given expiration 				date preceeds the 90 days, then the given expiration must be honored.

2. Accu-Chek Inform II strips are stable to the expiration date on the test 				strip bottle. The test strips must be used immediately after being removed from 			the bottle. Shut and seal the strip bottle immediately after removing strips.

3. No “open” dates are necessary on either controls or strips, but the control 			material MUST have the written expiration date clearly indicating that it is 			the expiration date. Use “exp.” or an “X” before the date.

C. FREQUENCY

1. Control Levels 1 (low) and 2 (high) will be run once in a 24 hour period. The 			meter will lock-out further patient testing until QC is performed and an acceptable 		result is obtained.

2. QC will also be performed in the following situations:

a. Each time a new vial of test strips is opened

b. When a vial of strips has been left opened more than 60 seconds

c. If the Accu-Chek Inform II™ System has been dropped

d. When test results contradict clinical symptoms

e. If questionable test results are displayed repeatedly

f. If you wish to test the performance of the sytem.

3. Patient testing may only proceed when quality control results are within the 			acceptable control range. This is indicated as PASS or FAIL. If the QC results 			FAIL, the problem must be corrected before any patient testing is performed.

D. CORRECTIVE ACTION FOR QC FAILURE

1. Verify that the strips and control material are not expired.

2. Ensure that the correct level is being tested. Repeat QC, if the result is still 			“FAIL” continue to step 3.

3. Run QC with a new bottle of strips, continue to step 4 if not resolved.

4. Run QC with new control solutions.

5. If QC is still failing, contact the Ancillary Testing Coordinator in the laboratory 		at ext. 3251

E. QC REVIEW

1. The Ancillary Testing Coordinator (LAB) will review QC on a monthly basis, 	and notify the Section Manager when corrective action is needed for identified 	problems.

2. The ATC will oversee performance quality related to glucose testing. This 	includes checking dating of control solutions, operator 	certification/recertification, review of QC results and monthly error reports, as 	well as acting on QC results which identify trends that may indicate potential 	problems.

F. QUALITY CONTROL PROCEDURE

1. Put on disposable gloves.

2. Press power ON button.

3. Enter your operator ID (last 4) and then press the forward arrow button.

4. Select Control Test.

5. Press the Barcode button and scan the bar code for either one of the control 			solutions bottles:  Level 1 (Low) or Level 2 (High).   Ensure the solution has not 			expired either by viewing manufacturer’s expiration date on new bottles or the 			hand written expiration date on previously open bottles (90days from date first 			opened).

6. Scan the test strip vial barcode. Check manufacturer expiration date on bottle 	and discard if expired.

7. Remove a test strip from the vial and replace the vial cap immediately.

8. When the flashing strip icon appears on the meter display, gently insert test 			strip with the yellow target area or test window facing up and insert the end with 			the gold bars.

9. When prompted, touch a bead of the control solution to the end of the test strip 			while holding or placing the meter horizontally. Do not hold the meter upright to 			run controls.

10. If you receive a “PASS” on both levels of control, proceed with patient 			testing. If QC receives a “FAIL,” choose a comment from below and proceed 			with corrective action. See section D above.

Quality Control Comments

Operator Error

Invalid, Will Repeat

Switched QC vial

Replaced test strips

Replaced QC solution

V. SAMPLE/SPECIMEN

Refer to the test strip package insert for the most current information

A. Capillary, venous and arterial whole blood specimens may be used for testing on 	the Accu-Chek Inform II™ System with Accu-Chek Inform II test strips. Do not use 	serum or plasma.

B. Point of care testing staff will use only the Capillary specimen.

C.  The capillary fingertip sample must be tested immediately after collection. Sufficient 	sample size is required to ensure accurate results. Do not “Milk” fingers for sample.

VI. PATIENT TESTING PROCEDURE

1. Put on disposable gloves.

2. Identify the patient. The operator must view the patient’s ID Badge and verify Date of Birth or Social Security Number. (The Joint Commission National Patient Safety Goal, the use of two patient identifiers).

3. Cleanse the puncture site (the side of the fingertip) with an alcohol swab and allow it to 	thoroughly dry. Alcohol at the puncture site must be dry or an error code/inaccurate result 	may occur.

4. Turn on the meter and enter your operator ID.

5. Choose Patient test from the Main Menu

6. Enter the patient identification by barcode scanning of the patient’s name badge.

Scan patient by bar code scanner function is used to enter patient ID in the Accu-Chek Inform II™ System. Manual input should be used only in situations where scanning is not available or not feasible.

7. You will now see a picture of a test strip with a downward flashing arrow on the screen indicating that you are ready to insert a test strip into the meter.

8. Remove a test strip from the vial and immediately recap the vial. Insert the test strip 	into the meter. The meter will display a flashing drop above the test strip icon when the 	test strip is properly inserted indicating that you are ready to apply a blood sample.

9. Perform a fingerstick using a single use lancet device. With a cotton ball, wipe the first drop of blood from the finger.

10. Apply next drop of blood to the front edge of the test strip. The sample will fill the yellow sample chamber by capillary action. Do not apply sample to the top of the test strip.

11. Once sufficient sample has been detected, the measurement begins. An hourglass icon 	indicates that the measurement is in progress.

12.After the sample has been obtained, apply gentle pressure to the puncture with a clean 	gauze square or cotton ball site for several minutes. If the patient is conscious and 	capable, enlist the patient’s assistance with applying pressure.

13.The measurement is complete when the result is displayed on the screen. Depending upon how high or low the result is, it may appear in a numeric or non-numeric format.  See Interpretation of Results section below for interpretation of each result format.

14. Remove the test strip and dispose of it in the trash receptacle.

15. Clean the meter with the facility approved antimicrobial wipes, ensuring no liquid 	gets in the strip reservoir. Cleaning must take place after each patient.

16. Dry the metal contacts and place the meter back on the docking station.

VII.	INTERPRETATION OF RESULTS

A. EXPECTED VALUES

1. The normal fasting blood glucose range in this facility for a non-diabetic adult 			is 71-109 mg/dL.

2. One to two hours after meals, normal blood glucose levels for a non-diabetic 			adult should be less than 145 mg/dL. The recommended level for a diabetic 			patient is a peak postprandial glucose less than 180 mg/dL.  Peak level blood 			glucose values depend on many factors including: age, stress, current overall 			control of blood sugars, insulin timing/administration, and medications.

B. CRITICAL RESULTS

1. A critical result is less than 40 mg/dL or greater than 400 mg/dL.

2. The following steps MUST be taken when dealing with a critical value:

1. Mark first result as Critical Result, RN Notified.

2. All critical results WILL be repeated immediately by an RN.

3. The repeated critical value MUST have the “MD Notified” comment. 				All RN repeats for critical values need this comment, even if it is the same 			patient on the same day. Only charting this in a note in the patient’s chart 				is not sufficient.

C. RESULTS OUTSIDE REPORTABLE RANGE

1. The Accu-Chek Inform II meters have a validated reportable range from 35 			mg/dL to 545 mg/dL.

2. This means that any result below 35 or greater than 545 will be reported in the 			following manner: <35 mg/dL or >545 mg/dL.

3. The meters are programmed to give a flag of RRHI or RRLO in the event the 			result is outside the verified range. If you receive an actual value, DO NOT chart 			the value, follow step 2 above.

4. It is recommended that a glucose test be performed by the laboratory via 			venipuncture sample to verify how high or low the patient’s glucose 				value truly is.

D. REPORTING

1. All valid results on the meters will be uploaded to RALS when the meter is 			docked. RALS will then upload results into the patient’s chart via VISTA.

2. The results will be delayed in RALS if results do not have the proper comments 		attached or have invalid comments attached.

3. Verification of results crossing over to the patient’s chart from the meters needs 		to occur at the end of each shift. This can be achieved by checking the laboratory 			reports in VISTA or CPRS.

4. In the event of downtime related to network failure or inability to upload data 			to CPRS/VISTA, glucose testing information may be recalled from each meter for 		up to 10 days or 4000 tests, whichever occurs first.

5. If the results do not appear to be crossing over, contact the Ancillary Testing 			Coordinator in the lab via email or phone. Ext. 3251

VIII. LIMITATIONS

## A. Test strips give dependable test results if the following limitations are understood:

1. The Accu-CHEK Inform II test strips are for testing fresh capillary, venous, 	arterial, or neonatal whole blood.  Cord blood samples cannot be used.

2. Hematocrit should be between 10-65%

3. Lipemic samples (triglycerides) in excess of 1800 mg/dL may produce elevated 	results.

4. Blood concentrations of galactose>15mg/dL will cause overestimation of blood 	glucose results.

5. Intravenous administration of ascorbic acid which results in blood 	concentrations of ascorbic acid >3 mg/dL will cause overestimation of blood 	glucose results.

6. If peripheral circulation is impaired, collection of capillary blood from the 	approved sample sites is not advised as the results might not be a true reflection of 	the phsiological blood glucose level.  This may apply to the following 	circumstances: severe dehydration as a result of diabetic detoacidosis or due to 	hyperglycemic hyperosmolar non-ketotic syndrome, hypotension, shock,  	decomepensated heart failure NYHA Class IV, or peripheral arterial occlusive 	disease.

7. This system has been tested at altitudes up to 10,000 feet.

8. The performance of this system has not been evaluated in the critically ill.

IX	. MAINTENANCE AND TROUBLESHOOTING

A. MAINTENANCE

1. Meters will be cleaned by the facility approved anti-microbial wipes 				after each use.

2. When cleaning with wipes, it is imperative that the strip port be covered 			so that no liquid will enter the port.

3. Wipe the metal contacts on the back of the meter dry before placing the 				meter on the docking station.

B. TROUBLESHOOTING

1. The following are the most common issues encountered with the 				glucose meters at this facility:

a. Buzzing sound coming from the docking station.

Resolution: The meter is not docked properly if there is a 						buzzing sound. Wiggle the meter and make sure it is properly 					seated. If the buzzing continues take a look at the metal contacts on 				the docking station itself. Sometimes the metal prongs become 					recessed and need to be pulled out slightly.

b. Meters not charging properly or not turning on at all.

Resolution: Check power supply to the meter docking station to 					make sure it hasn’t becom unplugged. Improper docking of the 					meter is usually the issue though. Make sure meter is properly 					docked. Clean contacts with a soft tissue and alcohol if needed, 					making sure to dry them. Check the prongs to see if they are 					recessed and pull them out slightly if they are.

c. Meters frozen or not very responsive to touch.

Resolution: This requires a power down. Press and hold the power 					button for about 20 seconds. The meter will shut off and cycle on. 					Once the meter comes back on you can try logging on or docking 					the station.

d. Invalid strip or invalid control error message.

Resolution: This can only be resolved by the Ancillary Testing 					Coordinator in the lab.

e. Inability to log on with user ID.

Resolution: Your certification has expired. Contact the Ancillary 					Testing Coordinator to become re-certified.

f. Results not crossing over to VISTA/CPRS.

Resolution: This is most likely due to the RALS server being 					offline. Contact the Ancillary Testing Coordinator.

g. Base units blinking red.

Resolution: Make sure the data ports are connected to the meter. 					They will also blink red when RALS is down. Contact Ancillary 					Testing Coordinator.

Note: If the above issues continue to be a problem or can’t be resolved you can 			contact the Ancillary Testing Coordinator in the lab at ext. 3251 or via email. The 			above issues are the most common. See appendix A for other issues and error 			alarms that may occur.

X. CERTIFICATION AND PROFICIENCY TESTING

A. CERTIFICATION

1.Only personnel whose certification and competency can be tracked in 				the Laboratory glucose database are permitted to perform patient testing 				independently. Agency nurses may perform CBG testing after obtaining 				computer access and CBG certification. Student nurses may perform CBG 			testing under their own access if they have computer (VistA) access and 				have been certified for CBG testing.

2.Operators must maintain certification to perform glucose testing and be 				proficient in the use of the Accu-Chek Inform II™ System. Employees 				authorized to perform glucose testing after completing the certification 				process include: RNs, LPNs, NAs, SNTs, MAs, Student Nurses, NPs, 				CNSs, MDs, and Health Technicians.

3.Initial certification and training records will be maintained in the 					laboratory.

4.To become a certified operator of the Accu-Chek Inform II™ System 				the first time, each operator will demonstrate the following:

a. Achievement of the knowledge and skills to perform blood 					glucose testing as defined in this policy/procedure. Initial users 					will be reassessed after 6 months.

b.Achieve a PASS on a complete Quality Control (QC) routine.

c.Be observed by a certified trainer performing a patient test which 			includes obtaining the specimen. (Employee testing of self or other 			employees for training is not acceptable.)

d.The meter will not allow a new operator access to perform 			testing until the certification information has been entered into the 			Lab Information Management System.

5.Recertification requires the following:

a.A passing score  on the knowledge-based  quiz.

b.Observation of testing performance by the Ancillary Testing 			Coordinator or certified trainer.

c.Perform a successful QC routine (both levels of controls) after 			the first six months of certification and at least annually thereafter. 			(Ongoing certification is automatically tracked via the RALS® 			System ).

d.Meters will not work for the operator after certification has 				expired.

6.Each operator will also be assessed for ongoing competency based on:

a. Monthly unit-based review of quality control and statistical 			reports.

b. Individualized findings from ongoing review of flagged results 			and errors by the Ancillary Care Testing Coordinator.

7. Retraining and reassessment will occur if problems are identified with 	user performance.

B. PROFICIENCY TESTING

1. Randomly selected operators will be requested to run tests on unknown 			samples according to the College of American Pathology (CAP) 				proficiency testing methodology to verify meter accuracy and operator 			competency. There will be no interdepartmental communication about 			proficiency testing samples until after the deadline for submission of the 			data to the proficiency testing provider. Referral of proficiency testing 			samples to another laboratory is prohibited.

2. Users are required to perform proficiency testing at least annually.

3. The test sample may be a blood product or derivative; therefore, 				standard precautions, including glove use must be observed.

X.	REFERENCES

College of American Pathologists. Point-of-Care-Testing Checklist. Northfield, 21 04 2014.

Roche Diagnostics. "Accu-Chek Inform II Package Insert." 2013.

Roche Diagnostics, Total Quality Management: Accu-Chek Inform II™ System: “Policies and Procedures and In-Service Program”. 2012. <Roche Diagnostics, Total Quality Management: Accu-Chek Inform II™ System: “Policies and Procedures and In-Service Program”>.

VETERANS AFFAIRS, PORTLAND. POC Accu-Chek Inform II Procedure Manual. n.d. PROQUIS.

Appendix A.

| Error Code/Message | Interpretation | Operator Action(s) |
| --- | --- | --- |
| LO/HI | Test result may be below (LO) or above (HI) than the reading range of the meter. | If result is unexpected based on patient’s condition, repeat the patient test with a new test strip. If patient test still LO or HI, report to an RN to rerun and verify the result and report to MD. QC should also be done on the meter. If control result is not within acceptable range, do not perform any further patient testing with that meter. The result will not be uploaded into the patient record until it has been evaluated. |
| RR LO/RR HI | Test result may be below/above the reportable range set by the system administrator. | If result is unexpected based on patient’s condition, repeat the patient test with a new test strip. If patient test still RRLO or RRHI, report to an RN to rerun and verify the result and report to MD. Confirmation glucose should be performed by the laboratory. QC should also be done on the meter. If control result is not within acceptable range, do not perform any further patient testing with that meter. The result will not be uploaded into the patient record until it has been evaluated. |
| CR LO/CR HI | Test result may be below/above the critical range set by the system administrator. | If result is unexpected based on patient’s condition, repeat the patient test with a new test strip. If patient test still CR LO or CR HI, report to an RN to rerun and verify the result and report to MD. |
| Strip Defect Error | Test strip may be damaged or the test result is extremely low and below the meter’s measurement range. | The test strip should be inserted into the meter prior to applying blood to the test strip. If this display appears before blood is placed on the strip, remove the test strip and reinsert. If the error display remains, repeat the test with a new strip. |
| Type Bad Dose | Insufficient amount of blood on the test strip. | Discard test strip and repeat the test. If the error persists, contact the ATC at 3251. |
| (QC) Fail or Out of Range | Failed QC test as it was out of range for the test strip lot. | Repeat control test with a new test strip. If the error persists, contact the ATC at 3251. |
| Glucose Error | Detection of an unexpected hardware error | Repeat the test, Run a control test with a new test strip. Reset the meter by following steps in IX. Maintenance and Troubleshooting section B part c. If the error persists, contact the ATC at 3251. |
| Unexpected SW Error | Detection of an unexpected software error | Repeat the test. Place the meter into a connected base unit to synchronize. Reset the meter by following steps in IX. Maintenance and Troubleshooting section B part c. If the error persists, contact the ATC at 3251. |