Beckman Coulter UniCel DXC 800

System Calibration

Policy		The analyzer must have a valid calibration in memory before controls and

patients can be processed.  The Modular Chemistry (MC) side must be

calibrated once every 24 hours and when a new modular chemistry reagent

is loaded.  Cartridge Chemistry reagents must be calibrated as indicated on the

Rgts/Cal icon. Electrolytes reagents must be calibrated every 8 hours.

Workplace	All laboratory employees are expected to maintain a safe working environment

Safety		and an injury-free workplace.  Laboratory employees are responsible for their

own safety, and the safety of others and adhering to all departmental and medical

center safety policies and procedures.

- For standard precautions and safety practices in the laboratory;  see LGM

8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids.

- For proper hand-washing, see LGM 8010, specifically, not limited to, proper hand-washing.

- For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves.

- For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste.

- For proper cleaning of work area, see LGM 8007  –  Cleaning Work Areas.

- For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- For proper storage and disposal of chemical hazardous waste, see LGM 8012.

Beckman Coulter UniCel DXC 800

System Calibration

Reagents		Aqua Calibrators 1,2,and 3

And			Protein Calibrator

Materials		Multi Calibrator

Alcohol Calibrator

Ammonia Calibrator

Bilirubin Calibrator

Cal 5 Plus Calibrator

Drug Calibrator 1 (Carbamazepine, Phenobarb, Theophylline, Phenytoin)

Drug Calibrator 2 ( Digoxin, Acetaminophen)

Drug Calibrator 3 Plus (Gentamicin, Tobramycin)

Enzyme Validator

Salicylate Calibrator

Vancomycin Calibrator

M-TP Calibrator

Thermo Lithium Calibrator

Data-Cal Chemistry Calibrator (Albumin, BCG)

Procedure		Load a Calibrator Disk

| Step | Action |
| --- | --- |
| 1. | Make sure the DXC I on Standby |
| 2. | Select Rgts/Cal from the menu bar |
| 3. | Select Option F6 |
| 4. | Select Load Calibrator diskette |
| 5. | Verify the lot number on the diskette is the same as the lot number on the calibrator bottle. Insert the diskette into the disk drive. Select OK |
| 6. | Verify that the lot number on the screen is the same as the lot number on the calibrator bottle |
| 7. | Remove the diskette from the disk drive - If no other diskettes are to be loaded select Cancel - If additional calibrator diskettes need to be loaded, , select OK and repeat Step 5 - If the incorrect diskette was placed in the disk drive, select OK, obtain the correct diskette and repeat Step 5 |

Beckman Coulter UniCel DXC 800

System Calibration

Procedure		Calibrator Assignment

| Step | Action |
| --- | --- |
| 1. | Select Rgts/Cal from the menu bar |
| 2. | Select Assign F7 |
| 3. | Use the down arrow option button to open lists of calibrators |
| 4. | Select the calibrator |
| 5. | Enter the ID of the calibrator or the rack/position number - If using a bar code label, the label must match the assigned calibrator ID - If a rack/position is assigned, the calibrator must be placed in this assigned position |
| 6. | Repeat Steps 3-5 to assign additional calibrator |
| 7. | When calibrator assignment is complete, select Close |

Calibrate DXC Chemistries with Bar Code Labels

| Step | Action |
| --- | --- |
| 1. | Select Rgts/Cal from the menu bar |
| 2. | Select the chemistries to be calibrated. If necessary use the up and down arrows to select more chemistries |
| 3. | Select Cal F4 |
| 4. | Select List F5 to review the calibrator list. Make sure the lot number on the calibrator bottle agrees with the lot number seen on the screen |
| 5. | Place a calibrator bar code label (with the name of the calibrator to be used) on a 16 mm transport tube. Repeat for all calibrators. Make sure the correct bar code labels are in the correct position |
| 6. | Place calibrator samples into sample cups - For Aqua Cals 1,2, and 3, fill a 0.5ml cup - All other calibrators need 3 drops of calibrator in a 0.5ml cup |
| 7. | Place the rack in the autoloader with the rack bar code label to the right. Press Run on the analyzer |

Beckman Coulter UniCel DXC 800

System Calibration

Procedure		Calibrate DXC Chemistries Using Assigned Calibrator Racks

| Step | Action |
| --- | --- |
| 1. | Select Rgts/Cal from the menu bar |
| 2. | Select the chemistries to be calibrated. If necessary, use the up and down arrows to select more chemistries |
| 3. | Select Cl F4 |
| 4. | Select List F5 to review the calibrator load list. Make sure the lot number on the calibrator bottle is the same as the lot number on the screen |
| 5. | Get the assigned calibrator racks listed on the screen and place the cups into assigned rack positions as shown on the load list |
| 6. | Place calibrators into sample cup Place calibrator samples into sample cups - For Aqua Cals 1,2, and 3, fill a 0.5ml cup - All other calibrators need 3 drops of calibrator in a 0.5ml cup |

Cancel a Calibration Request

| Step | Action |
| --- | --- |
| 1. | Select Rgts/Cal from the menu bar |
| 2. | Select the appropriate requested and assigned chemistries to be cancelled |
| 3. | Select No Cal F8 to cancel calibration request Chemistries currently selected and programmed for calibration may be cancelled only if the calibration is not yet in progress. Calibration reports for a given calibrator do not print until all calibrations are complete for each chemistry requested, unlike sample and control results, which print on a cup-by-cup basis |

Beckman Coulter UniCel DXC 800

System Calibration

Procedure		Enzyme Validator

| Step | Action |
| --- | --- |
| 1. | The Enzyme Validator option provides the ability to obtain IFCC/DGKCh equivalent answers to selected enzymes, or to obtain IFCC DGKCh equivalent answers for the non-IFCC/DGKCh formulation enzymes. When the enzyme reagent is configured and loaded, and the Enzyme Validator feature turns on, Enzyme Validator assignments and calibration requests can be performed in the same manner as any calibrator Select Rgts/Cal from the menu bar |
| 2. | Select Options F6 |
| 3. | Select Enzyme Validator |
| 4. | Select a desired Configured Chemistry name from the pull down menu |
| 5. | Select Add to move the chosen Chemistry to the Selected Chemistries area |
| 6. | Repeat Step 5 for other chemistries, if appropriate |
| 7. | Select OK to enable the Enzyme Validator option OR Select Cancel to exit without enabling the option Additional Information: DGKCh - German Clinical Chemistry Association IFCC (International Federation of Clinical Chemistry), a committee that establishes formulation standards in Europe |

Beckman Coulter UniCel DXC 800

System Calibration

Procedure		Calibration Override

| Step | Action |
| --- | --- |
| 1. | Overriding a calibration is not usually justified. However, there may be an emergency situation when a delay caused by calibration is unacceptable to the laboratory. The magnitude or error which is deemed when overriding a failed calibration is totally under the control of the laboratory. Therefore, the Override function should be used with caution Select Rgts/Cal from the menu bar |
| 2. | Select the Chem position to override |
| 3. | Select Options F6 |
| 4. | Select Calibration Override from the pull-down menu The Cal Overridden status can only be removed by successfully recalibrating the chemistry or by selecting Chem bypass. A flag indicating that a particular chemistry has been overridden will appear in the Instrument code area on the appropriate Summary or Patient Chartable Results report. Although a Cal Override status can be bypassed, the Cal Overridden flag is only removed by a successful calibration |

Chemistry Bypass

| Step | Action |
| --- | --- |
| 1. | The system will not aspirate sample or dispense reagents for bypassed chemistries. They will not be marked as Pending. Bypassed chemistries are logged on the Post-Run Summary Report. Bypassed chemistries must be calibrated to process samples Select Rgts/Cal from the menu bar |
| 2. | Select the Chem position to bypass |
| 3. | Select Options F6 |
| 4. | Select Bypass Chemistry from the pull-down menu To remove a Bypassed Status: Recalibrate the chemistry to remove the bypassed status |

Beckman Coulter UniCel DXC 800

Reagent and Material Management

Controlled		The controlled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
| LGM 8000 | Safety Practices |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |
| LGM 8006 | Handling of Regular and Infectious Waste |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand-washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |

Non			The non controlled documents are as follows:

Controlled

Documents

| Document No. | Document Name |
| --- | --- |
|  | Operator’s Manual |