#### ACCUMETRICS VERIFYNOW MAINTENANCE FOR P2Y12 TEST

| Policy | This section contains the necessary information to perform the scheduled Preventive Maintenance of the Accumetrics VerifyNow. The operators must follow this schedule to ensure years of reliable performance, minimal downtime, and to prevent costly factory and field services repairs. |
| --- | --- |

| ##### Safety or Special Safety Precautions | All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety and the safety of others and adhering to all departmental and medical center safety policies and procedures. - For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). - For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. - For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing. - For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. - For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste. - For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. - For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. - For proper storage and disposal of chemical hazardous waste, see LGM 8012. |
| --- | --- |

- Continued on next page

#### ACCUMETRICS VERIFYNOW MAINTENANCE FOR P2Y12 TEST

| ##### Materials and supplies |  | - EQC - Preventive Maintenance kit/ Cleaning devices - Fan filter - Lint free cloth ( clean room wipes) or cotton swab - Non-abrasive cleaner (such as 10% household bleach solution in water, dish washing detergent or isopropyl alcohol) |  |  |
| --- | --- | --- | --- | --- |
| ##### Procedure: | - Daily Maintenance - Run Electronic Quality Control (EQC) - Refer to LCM 287 for procedure on how to run EQC. - Bi-weekly Maintenance: Use the Cleaning Device |  |  |  |
|  | - Depending on the room conditions, small amounts of dust and debris may build up on the pneumatic port connection of the instrument. A single use, disposable cleaning device is provided to remove this debris. It consists of a clear plastic component with an adhesive strip. Use the cleaning device to remove debris from a pneumatic port connection inside the test device port according to the following procedure; |  |  |  |
|  | Step |  | Action |  |
|  | 1 |  | Locate the cleaning device pack. |  |
|  | 2 |  | Grasp the cleaning device using the finger grip and remove the clear plastic tape. Discard the tape. |  |
|  | 3 |  | Open the cover and insert the cleaning device into the test device port until it clicks. The adhesive strip removes dust and debris on the internal pneumatic port cleaning strip. NOTE: The instrument does not need to be in any particular mode to use the cleaning device. |  |
|  | 4 |  | Leave the cleaning device in the test port for five seconds, but never longer than ten seconds. CAUTION: Leaving the cleaning device in the test port for an extended period of time can damage the instrument from adhesive adhering to the internal pneumatic port. |  |
|  | 5 |  | Remove the cleaning device completely from the test device port. Discard the device. The cleaning device is for single use only. |  |
|  | 6 |  | Repeat using another cleaning device if, upon inspection, there is visible dust and debris. |  |

#### ACCUMETRICS VERIFYNOW MAINTENANCE FOR P2Y12 TEST

|  | - Monthly Maintenance: Clean Exterior Surfaces |  |
| --- | --- | --- |
|  | - Occasionally, a liquid may spill near the instrument or small amounts of dust may build up on the exterior of the instrument. Periodic cleaning of the exterior surface may be performed to remove accumulation on the instrument. |  |
|  | Step | Action |
|  | 1 | Cleaning solutions may be applied by moistening a soft, lint-free cloth, and/or a cotton swab. Do not allow liquids to flow freely or be sprayed on the instrument |
|  |  | NOTE: Use the recommended cleaning solution on the exterior surfaces. |
|  | 2 | After cleaning with one of the solutions, a cloth moistened with fresh water should be used to dilute and remove all of the residual cleaning solution from the instrument’s surfaces. |
|  | 3 | Wipe the EQC device with a damp cloth moistened with isopropyl alcohol. |
|  | - Annual Maintenance: Replace the Fan Filter |  |
|  | - The instrument is equipped with an exterior fan vent to cool the internal electronic components. Depending on the room conditions, small amounts of dust and debris may accumulate over time within the fan filter and obstruct airflow. Periodic cleaning of the fan filter may be performed to avoid overheating of the instrument. |  |
|  | Step | Action |
|  | 1 | Power off the instrument (“O” designates off). |
|  | 2 | The fan filter is held in place by a press fit retainer. Using your fingernail or a small tool, gently pull from the center of the filter retainer. NOTE: There are no screws to be removed. |
|  | 3 | Remove the fan filter and inspect it for dust accumulation. |
|  | 4 | If necessary, replace it with a new filter. |
|  | 5 | Gently replace the fan filter and plastic filter retainer. Do not position the instrument so that airflow to the fan is obstructed. |
|  | 6 | Power on the instrument (“I” designates on). |

- NOTE: Document all maintenance procedure done in the Maintenance Log Sheet.

#### ACCUMETRICS VERIFYNOW MAINTENANCE FOR P2Y12 TEST

| ##### Non-Controlled Documents | The following non-controlled documents support this procedure. - Operator’s Manual |  |
| --- | --- | --- |
| ##### Controlled Documents | The following controlled documents support this procedure. |  |
|  | Procedure | Number |
|  | Accumetrics VerifyNow Sample Processing for P2Y12 test | LCM 286 |
|  | Accumetrics VerifyNow Quality Control for P2Y12 test | LCM 287 |
|  | Safety Policies | LGM 8000 |
|  | Infection Control | LGM 8004 |
|  | Universal Body Substance Precaution | LGM 8005 |
|  | Handling of Regular and Infectious Waste | LGM 8006 |
|  | Cleaning Work Areas | LGM 8007 |
|  | Hand washing Policy | LGM 8010 |
|  | Storage and Disposal of Chemical Hazardous Waste | LGM 8012 |

| ##### Author Aida R. Legaspi, CLS |  |
| --- | --- |

#### ACCUMETRICS VERIFYNOW MAINTENANCE FOR P2Y12 TEST

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Aida Legaspi, CLS Area Laboratory Manager – Los Angeles Medical Center |  |
| Julie Toti, CLS, MS Laboratory Director – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

#### ACCUMETRICS VERIFYNOW MAINTENANCE FOR P2Y12 TEST

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |