#### PROFILE®-V MEDTOX SCAN® SAMPLE PROCESSING FOR DRUGS OF ABUSE TEST SYSTEM

| Purpose | The Profile®-V MEDTOX Scan® Drugs of Abuse test system is used for the qualitative detection of the drugs of abuse. ____________________________________________________________ |
| --- | --- |
| Scope Policy | The Profile®-V MedTox Scan® Drugs of Abuse Test System consists of the Profile®-V MedTox Scan® test Devices and the MedTox Scan Reader. The Profile®-V MedTox Scan® Test Devices are one-step immunochromatographic tests for rapid, qualitative detection of one or more of the following in human urine: Amphetamine, Barbiturates, Benzodiazepines, Buprenorphine, Cocaine, Methadone, Methamphetamine, Opiates, Oxycodone, Phencyclidine, Propoxyphene, THC (Cannabinoids) and Tricyclic Antidepressants or their metabolites. ____________________________________________________________ The Profile®-V MedTox Scan® Test Devices can only be used with the MedTox Scan Reader. The MedTox Scan Reader is an instrument used to interpret and report results of the Profile®-V MedTox Scan® Test Devices. The MedTox Scan Test devices cannot be visually read. |

Specimen         The urine sample should be collected in a clean, dry container.

Collection         Approximately 75µL is required for each sample well. Collection of 30 ml

urine is more than sufficient for initial and subsequent testing. No

preservatives should be added. Urine may be tested immediately

following collection.

____________________________________________________________

Specimen         If it is necessary to store the urine, store under refrigeration at 2 to 8°C (36

Storage             46°F) for no more than two days. Urine may be frozen at -20°C (-4°F).

Stored urine must be brought to ambient temperature (18-25°C/64-77°F)

and mixed well to assure a homogenous sample prior to testing.

#### _____________________________________________

#### PROFILE®-V MEDTOX SCAN® SAMPLE PROCESSING FOR DRUGS OF ABUSE TEST SYSTEM

| Specimen rejection | - Urine collected in a container with preservatives |
| --- | --- |

| ##### Reagents and/or Media | Description | Vendor | Storage |
| --- | --- | --- | --- |
|  | Profile®-V MEDTOX Scan® Drugs of abuse test system kit | MEDTOX Diagnostics | Room temperature |

| ##### Materials and supplies | - Disposable pipette tips - MiniPet pipettor - Thermal printer and printer paper |
| --- | --- |

- Continued on next page

#### PROFILE®-V MEDTOX SCAN® SAMPLE PROCESSING FOR DRUGS OF ABUSE TEST SYSTEM

___________________________________________________________

| ##### Safety or Special Safety Precautions | All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety and the safety of others and adhering to all departmental and medical center safety policies and procedures. - For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). - For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. - For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing. - For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. - For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste. - For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. - For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. - For proper storage and disposal of chemical hazardous waste, see LGM 8012. . |
| --- | --- |

| ##### Definitions | The Profile®-V MEDTOX Scan® Drugs of Abuse Test System is a single use, one-step, rapid immunoassay device that gives the Clinical Laboratory the ability to provide stat toxicology assays. |
| --- | --- |

- Continued on next page

#### PROFILE®-V MEDTOX SCAN® SAMPLE PROCESSING FOR DRUGS OF ABUSE TEST SYSTEM

| ##### Procedure: | - Follow the steps below to run patients…. |  |  |
| --- | --- | --- | --- |
|  | Step | Action |  |
|  | 1 | Open one pouch for each sample to be tested and mark the Profile®-V MedTox Scan Test Device with the patient or sample identification (ID). Make sure you only mark along the left edge of the test device (labeled “ID >). You may notice a reddish-purple color in the sample well. This is normal, do not discard the test. |  |
|  | 2 | Dispense 75µL of urine into sample well (indicated by ▼on the test device) - Place a disposable yellow sample tip securely onto the end of the green (75µL) MiniPet. - Grasp the MiniPet under its collar using the index and middle fingers. With the thumb, depress the plunger completely. - Holding the MiniPet vertically (straight up-down), lower the yellow tip no more than ¼ “ into the urine specimen. - With tip in the urine specimen slowly and smoothly release the plunger allowing it to rise completely. - Visually inspect the urine sample in the tip. Ensure there are no air bubbles and that no excess urine is on the outer surface of the tip. - Hold the pipette tip directly over the sample well. Depress the plunger completely to dispense the entire contents of urine into one sample well of the testing device. |  |
|  | 3 | Repeat Step 2 for all sample wells with a ▼above them. |  |
|  | 4 | Place the test device in the MedTox Scan® Reader cassette drawer and close the drawer immediately. The MedTox Scan® Reader will read the barcode on the test device and determine its part number and test configuration. It will prompt the user to enter Lot#, User ID#, and Specimen ID#, which can all be entered using the MedTox Scan® Reader keypad or hand held barcode scanner. The MedTox Scan® Reader will begin timing the assay once it detects the barcode and results will be displayed after the scan and analysis are complete. |  |
|  | 5 | Discard disposable yellow MiniPet sample tip. Store the MiniPet in a dry, secure location at room temperature (18-25°C or 64-77°F). Replace the MiniPet if it becomes damaged or does not function properly. |  |

#### PROFILE®-V MEDTOX SCAN® SAMPLE PROCESSING FOR DRUGS OF ABUSE TEST SYSTEM

___________________________________________________________

| Interpretation/ Result | The MEDTOX Scan Reader will automatically read the control and test lines at the correct positions and display the test results for each drug. Print results and file. The MEDTOX Scan Reader displays the results as either “NEG” for a negative result, “POS” for a preliminary positive result, or “INVALID” for an invalid result. “VALID” will be displayed if valid results are obtained. Profile®-V MEDTOX Scan Test devices cannot be visually read. Valid: The control line must be present for the test to be valid. NEG: A Negative test result for a specific drug indicates that the sample does not contain the drug/drug metabolites above the cutoff level. POS: A preliminary Positive test result for a specific drug indicates that the sample may contain drug/drug metabolites near or above the cutoff level. It does not indicate the level of intoxication or the specific concentration of drug in the urine sample. Invalid: The control line must be present for the test to be valid. The absence of a control line indicates the test is invalid. The urine samples should be retested on a new test device. Confirmation of Positive Urine Toxicology Screen Results: All positive drug screens (except TCA) from children <14 yrs. old and from L/D patient will be confirmed as follows Drugs Detected; Confirmation sent to:. AMP, mAMP; Regional Reference Lab. PCP; Regional Reference Lab. BAR; QUEST. BZO; QUEST. COC; Regional Reference Lab. OPI; Regional Reference Lab. THC; Regional Reference Lab. For all other adults: Drugs Detected; Confirmation sent to:. AMP, mAMP; Regional Reference Lab. PCP; Regional Reference Lab. BAR; Do not confirm unless requested by Provider. BZO; Do not confirm unless requested by Provider. COC; Do not confirm unless requested by Provider. OPI; Do not confirm unless requested by Provider. THC; Do not confirm unless requested by Provider. All positive urines should be saved for 7 days. |
| --- | --- |

#### PROFILE®-V MEDTOX SCAN® SAMPLE PROCESSING FOR DRUGS OF ABUSE TEST SYSTEM

___________________________________________________________

| ##### Limitations |  | - The Profile®-V MedTox Scan® Drugs of Abuse Test System is only for use with unadulterated preservative free, human urine samples. - A presumptive positive result for any drug does not indicate the level of intoxication, administration route or concentration of that drug in the urine specimen. - A Negative result may not necessarily indicate drug-free urine. Negative results can be obtained when drug is present but below the cut-off level of the test. - Once the test device has been read, it must not be reinserted for repeat reading. Use a fresh test cassette if a repeat reading is required. - The Profile®-V MedTox Scan Drugs of Abuse Test System cannot distinguish between abused drugs and certain prescribed medications. |
| --- | --- | --- |
| ##### Non-Controlled Documents |  | The following non-controlled documents support this procedure. - Operators Manual |

| ##### Controlled Documents | The following controlled documents support this procedure. |  |
| --- | --- | --- |
|  | Procedure | Number |
|  | Safety Practices | LGM 8000 |
|  | Infection Control | LGM 8004 |
|  | Universal Body Substance Precaution | LGM 8005 |
|  | Handling of Regular and Infectious Waste | LGM 8006 |
|  | Cleaning Work Areas | LGM 8007 |
|  | Hand washing Policy | LGM 8010 |
|  | Storage and Disposal of Chemical Hazardous Waste | LGM 8012 |

| ##### Author Aida R. Legaspi, CLS |  |
| --- | --- |

#### PROFILE®-V MEDTOX SCAN® SAMPLE PROCESSING FOR DRUGS OF ABUSE TEST SYSTEM

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Aida Legaspi, CLS Area Laboratory Manager – Los Angeles Medical Center |  |
| Julie Toti, CLS, MS Laboratory Director – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

#### PROFILE®-V MEDTOX SCAN® SAMPLE PROCESSING FOR DRUGS OF ABUSE TEST SYSTEM

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |