#### PROFILE®-V MEDTOX SCAN® QUALITY CONTROL FOR DRUGS OF ABUSE TEST SYSTEM

| Purpose | The purpose of quality control is to ensure accuracy and reliability of results and to detect errors. ____________________________________________________________ |
| --- | --- |
| Scope Policy | To monitor the performance of the Profile®-V MedTox Scan® Test Devices and the MedTox Scan Reader a combination of internal controls and external controls must be done. Internal controls ensure that the test is working and that you are performing the test correctly. A control line (internal control) is included on each Profile®-V MedTox Scan® test strip. Whether or not drug is present in the sample, a line must form at the control position on the strip.to show that enough sample volume was used and that the reagents are migrating properly. External controls are urine-based control materials that contain the drugs to be tested at concentrations above the cutoff (positive control) or contain no drugs (negative control). ____________________________________________________________ Run external controls once a week in the same manner as the patient samples. Running controls should be done on a monthly rotation basis between 3 shifts. Run the controls also when you open a new lot of devices, if you suspect that the reader or test device is not working properly, if you have had a repeated unexpected test result, or if you suspect that the test device have been stored improperly. |

#### PROFILE®-V MEDTOX SCAN® QUALITY CONTROL FOR DRUGS OF ABUSE TEST SYSTEM

| Storage Stability | Unopened: - The controls are stable until the expiration date when stored at -10 to -20°C and protected from light. - The controls are stable until the expiration date when stored at 2-8°C. After Opening: - The controls are stable for 6 months or until the expiration date, whichever comes first, when stored at -10 to -20°C. (Controls can be aliquoted and frozen) - The controls are stable for 31 days or until the expiration date, whichever comes first, when stored tightly capped at 2-8°C. - Thaw controls as needed; allow controls to come to room temperature followed by gentle swirling before use. |
| --- | --- |

| ##### Materials and supplies | - Profile®-V MEDTOX Scan ® Drugs of Abuse Test System kit - MEDTOX Scan Positive and Negative Toxicology Urine Controls - Disposable pipette tips - MiniPet pipettor - Thermal printer and printer paper |
| --- | --- |

#### PROFILE®-V MEDTOX SCAN® QUALITY CONTROL FOR DRUGS OF ABUSE TEST SYSTEM

___________________________________________________________

| ##### Safety or Special Safety Precautions | All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety and the safety of others and adhering to all departmental and medical center safety policies and procedures. - For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). - For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. - For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing. - For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. - For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste. - For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. - For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. - For proper storage and disposal of chemical hazardous waste, see LGM 8012. . |
| --- | --- |

| ##### Procedure | Allow controls to come to room temperature followed by gentle swirling or inversion before use. Do not shake. Pipette an appropriate aliquot of MEDTOX control urine as required by the drugs of abuse test device or screening method. Controls should be treated as any “unknown” specimen while following the specific protocol of the assay being used. |
| --- | --- |

- Continued on next page

#### PROFILE®-V MEDTOX SCAN® QUALITY CONTROL FOR DRUGS OF ABUSE TEST SYSTEM

| Interpretation/ Result |  | The positive MEDTOX control must test positive on the Profile®-V MedTox Scan® Drugs of Abuse Test System. The negative control must test negative. Document and keep all QC print outs in the QC log sheet. ________________________________________________________________ |  |  |
| --- | --- | --- | --- | --- |
| ##### Limitations |  |  | This control is meant to be used to validate the performance of immunoassay drug screening methods. This product is not meant to be used as a standard or calibrator. The MEDTOX Scan Reader is calibrated and verified before shipment. No calibration is required of the operator. |  |

| ##### Non-Controlled Documents | The following non-controlled documents support this procedure. - Operators Manual |
| --- | --- |

| ##### Controlled Documents | The following controlled documents support this procedure. |  |
| --- | --- | --- |
|  | Procedure | Number |
|  | Profile®-V MedTox Scan® Sample Processing For Drugs of Abuse Test System | LCM 288 |
|  | Safety Practices | LGM 8000 |
|  | Infection Control | LGM 8004 |
|  | Universal Body Substance Precaution | LGM 8005 |
|  | Handling of Regular and Infectious Waste | LGM 8006 |
|  | Cleaning Work Areas | LGM 8007 |
|  | Hand washing Policy | LGM 8010 |
|  | Storage and Disposal of Chemical Hazardous Waste | LGM 8012 |

| ##### Author Aida R. Legaspi, CLS |  |
| --- | --- |

#### PROFILE®-V MEDTOX SCAN® QUALITY CONTROL FOR DRUGS OF ABUSE TEST SYSTEM

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Aida Legaspi, CLS Area Laboratory Manager – Los Angeles Medical Center |  |
| Julie Toti, CLS, MS Laboratory Director – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

#### PROFILE®-V MEDTOX SCAN® QUALITY CONTROL FOR DRUGS OF ABUSE TEST SYSTEM

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |