| # Purpose | To define the process of tracking apheresis donor’s total red blood cell and plasma loss over a 12 month rolling calendar to prevent these donors from exceeding maximum allowable red cell and plasma volume losses. |
| --- | --- |

| # Policy | - A donor must not undergo more than 24 Platelet Pheresis collections in a 12 month rolling period. - The interval between each collection of Platelet Pheresis must be at least two days with no more than two procedures in a seven-day period. - Concurrent plasma must not be collected more than every 28 days. - A donor must not undergo more than 12 concurrent plasma collections in a 12 month rolling period. - Donors must not donate red blood cells for 8 weeks/56 days after a single RBC apheresis or 16 weeks/112 days after a double RBC apheresis donation - Allowable Double Red cell donations per year is three (3) provided they do not exceed maximum RBC loss in rolling 12 month period - The automated blood cell separator device must be set with a post-donation platelet count target of no less than 100,000 platelets/µL. |
| --- | --- |
|  | - Maximum Allowable RBC Loss in a rolling twelve (12) month calendar is (1733 = 45% x 550 mL (488 mL draw + 62 mL sample pouch + tubing ) x 7). |
|  | - Cumulative plasma loss must not exceed 14.4 liters (14,400 mL) if the donor weighs 175 lbs. or more and 12.0 liters (12,000 mL) if the donor weighs 110- 174 lbs. in a rolling twelve month (12) period. |
|  | - Apheresis and whole blood donations must be tracked and red cell and plasma losses calculated at each donation. - Whole Blood donations need to be accounted for- Volume 550 mL- with 248 mL as RBC loss, 302 mL plasma loss (assumptions, 488 mL blood draw plus, 62 mL (sample pouch + tubing) collected and hct of 45% or hgb of 15.0 gm/dL). - All donations (including non KP donations) must have appropriate losses tracked (whole blood donations, platelet, plasma, single/double red cells, rinseback completed/no rinseback, etc.) and values documented in Volume Loss Worksheet. |
|  | Continued on next page |

|  | - All charts are reviewed by licensed staff prior to donation to determine eligibility based on volume loss and number of donations in a rolling calendar year. This review is documented on the Volume Loss Worksheet. - Donors who inadvertently exceed maximum allowable RBC loss must be deferred from any donations for the appropriate length of time. The chart is reviewed by the Donor Center Medical Director, QIM is completed, and notes documented on chart |
| --- | --- |
|  | - Donors who inadvertently exceed maximum allowable plasma loss or plasma is collected more than once in 28 days must be deferred for an appropriate period of time and an TP/SPEP performed. The chart is reviewed by the Donor Center Medical Director, QIM is completed, and notes documented on chart |

| Equipment and Reagents | - Volume Loss Worksheet (VLW) Attachment A - Procedure Runsheet- Attachment B - Calculator - Donor Apheresis Chart - Donor History Questionnaire (DHQ) - Computer System |
| --- | --- |

| Definitions | - RBC – Red Blood Cell - WB – Whole Blood - Hct – Hematocrit - DRBC – Double Red Blood Cell - TP/SPEP – Total Protein/ Serum Protein Electrophoresis - KP – Kaiser Permanente - YTD - year to date, rolling 12 month - DOB – date of birth - DID – donor ID number - AC – anticoagulant - VLW – volume loss worksheet – Attachment A - ERC- temporary open deferral code for donors who exceed red cell and/or plasma volume limits - NRB – temporary 56 day deferral code for donor when you have no rinseback. - QIM – Quality Improvement Monitoring (Tool) |
| --- | --- |
|  | Continued on next page |

| ## Procedure | Perform the following steps to complete the Volume Loss Worksheet (VLW) |  |
| --- | --- | --- |
|  | ## Step | Action |
|  | 1. | Document donor information at top of VLW (donor’s name, DID, DOB, height, weight, from DHQ)- Attachment A - Verify that the donor is eligible to donate based on the eligibility dates listed on the VLW. Place a checkmark next to the eligibility date to document the verification. |
|  | 2. | Complete the procedure information (date, unit number, hematocrit, platelet pre-count when obtained from lab, TP/SPEP if required). - TP/SPEP testing is required IF: - Plasma is collected more than once in 28 days. - Cumulative plasma volume loss exceeds maximum allowable loss in a rolling 12 month period - 14.4 Liters (14,400 mL) if donor weighs 175 lbs. or more and - 12.0 liters (12,000 mL) if donor weighs less than 175 lbs. - These calculations exclude anticoagulant volumes. - Document TP/SPEP in comments under the Product Yield column on VLW |
|  | 3. | Under the Procedure columns, document the procedure being performed, in the non-shaded section of the appropriate column(s). - For example if there is a single platelet and a plasma unit, mark “S” in the Platelet column and “S” in the Plasma column to designate a single platelet and single plasma. - A “D” in the Platelet column and an “S” in the RBC column would designate a double platelet and a single RBC. |
|  | 4. | Column labeled RBC Collected (column A): - Enter the volume collected from the “Trima Accel Automated Blood Component System Procedure Runsheet” (Runsheet) under Red Cell Product, RBC Collect Volume (mL) minus the anticoagulant (AC). - For example: 250 mL of RBC collected – 9 mL of AC = 241 mL (RBC Volume collected. Enter this result in column A of the VLW |
|  |  | Continued on next page |

| ## Procedure | Perform the following steps to complete the Volume Loss Worksheet (VLW) |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | ## Step | Action |  |  |  |
|  |  | Column labeled RBC Residual (column B): - RBC residual values require a series of calculations using the Runsheet. The value(s) are documented in appropriate columns for RBC Residual Loss (column B) on the VLW. See below and see attached sample VLW |  |  |  |
|  | Leukoreduced Red Cell Loss Calculation Loss Guide: |  |  |  |  |
|  | IF |  | THEN | EXAMPLE | WHERE TO DOCUMENT |
|  | Rinse-back Complete |  | See “End Run Summary screen” for RBC residual loss and plasma loss | Example: Hct = 45%; Hgb 15.0 gm/dL RBC Loss from test samples =62ml (pouch + sample tubing) RBC residual loss = 40ml (Refer to the Trima Procedural Summary Screen for actual RBC loss) (0.45 x 62mL) = 28ml (pRBC) 40ml(residual RBC)+28 (pRBC) = 68ml RBC loss | “Residual” column under RBC on theVLW, (column B) |
|  | No Rinse back and Blood Prime is complete |  | See “End Run Summary screen” for RBC residual loss and plasma loss | Example: Residual RBC volume loss = 85 mL ((Refer to the Trima Procedural Summary Screen for actual RBC loss) | “Residual” column under RBC on the VLW (column B) |
|  | Continued on next page |  |  |  |  |

| Procedure | Perform the following steps to complete the Volume Loss Worksheet (VLW) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | ## Step | Action |  |  |  |  |
| Cont’d | Leukoreduced Red Cell Loss Calculation Loss Guide: |  |  |  |  |  |
|  | IF |  | THEN |  | EXAMPLE | WHERE TO DOCUMENT |
|  |  |  |  |  | RBC Loss from test samples =62ml (pouch + sample tubing) (62 mL x Hct45%; Hgb 15.0 gm/dL)=28mlpRBC 85 mL(Trima Display) + 28ml pRBC(pouch+tubing) = 113 mL |  |
|  | No Rinseback and Blood Prime is not complete |  |  | See “End Run Summary screen” for RBC residual loss and plasma loss | Example: Residual RBC volume loss = 45 mL ((Refer to the Trima Procedural Summary Screen for actual RBC loss) RBC Loss from test samples =62 mL (pouch + sample tubing) (62 mL x Hct 45%; Hgb 15.0 gm/dL)=28 mL RBC 45 mL(Trima Display) + 28 mL pRBC(pouch+tubing) = 73 mL RBC Loss |  |
|  | Continued on next page |  |  |  |  |  |

| Procedure | Perform the following steps to complete the Volume Loss Worksheet (VLW) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | ## Step | Action |  |  |  |  |
| Cont’d | Leukoreduced Red Cell Loss Calculation Loss Guide: |  |  |  |  |  |
|  | IF |  | THEN |  | EXAMPLE | WHERE TO DOCUMENT |
|  | It becomes impossible to return donor’s red cells during apheresis AND RBC loss is 200 mL or greater |  |  | See “End Run Summary screen” for RBC residual loss plasma loss Note: At least 8 weeks shall elapse before a subsequent apheresis procedure, |  |  |
|  |  |  |  |  | Continued on next page |  |

|  | Leukoreduced Red Cell Loss Calculation Loss Guide: |  |  |  |
| --- | --- | --- | --- | --- |
|  | IF | THEN | EXAMPLE | WHERE TO DOCUMENT |
| Cont’d | It becomes impossible to return donor’s red cells during apheresis AND RBC loss is 300 mL or greater | See “End Run Summary screen” for RBC residual loss plasma loss Note: At least 16 weeks shall elapse before a subsequent apheresis procedure, |  |  |

| 5. | Column labeled Red Cell Loss this procedure/Red Cell loss year to date - “RBC Loss This Procedure”(column C) is calculated by adding RBC collected column (column A) to RBC residual column (column B) Column labeled: RBC Loss Year to Date (shaded) - RBC Loss YTD is calculated after each procedure and includes the volumes of all collections (column C) within the last 12 months. - Example- C1+C2+C3 (etc) = D; where C1, C2 etc are the current donation volume and the volumes from the most recent donations going back to 12 months. |
| --- | --- |
| 6. | Column labeled Platelets Collected: - Enter the volume collected from the Runsheet under Platelets Product (TRIMA), Platelet Product Volume (mL) minus the anticoagulant (AC) (AC is on the run sheet). - For example: 275mL of platelets collected – 40 mL of AC = 235 mL Platelet Volume collected. Enter this result in column E of the VLW |
| 7. | Column labeled Plasma Collected: - Enter the volume collected from the Runsheet under Plasma Product, Plasma Volume (mL) minus the anticoagulant (AC). - For example: 200mL of plasma collected – 30 mL of AC = 170 mL Plasma product collected. Enter this result in column F of the WLV Column labeled Plasma Residual (column G): |
|  | Continued on next page |

|  | Plasma Loss Calculation Loss Guide: |  |  |  |
| --- | --- | --- | --- | --- |
|  | IF | THEN | EXAMPLE | WHERE TO DOCUMENT |
|  | Procedure with Rinseback | See “End Run Summary screen” for residual plasma | See procedure run sheet | Plasma Residual column (column G) |
|  | Procedure with No Rinseback | See “End Run Summary screen” for residual plasma | See procedure run sheet | Plasma residual column (column G) |

| 8. | Column labeled Plasma Loss this procedure/Plasma loss year to date - “Plasma Loss This Procedure”(column H) is calculated by adding Platelet collected (column E) to Plasma collected (column F) to Plasma Residual (column G) - Plasma Loss Year to Date (column) is calculated after each procedure. A rolling 12 month calendar is used - Add the volume from the Plasma Loss this Procedure (column H) for each collection not to exceed a 12 month rolling calendar. - Example= Plasma loss YTD (column I) H+H+H(etc) = I |
| --- | --- |
| 9. | Column labeled Date of Donation/ Number of donation YTD - Document number of donations in the past 12 months in “# donations YTD” column - This requires you to go back to previous donations within a 12 month period and count the donations. |
| 10. | Columns labeled Eligibility Date (shaded) under each procedure - Document next eligibility date for Platelets, Plasma and RBCs in eligibility date columns on VLW, based on the following criteria: |
|  | Continued on next page |

|  | Apheresis Product | # donations in a 12 month period | Maximum loss in a 12 month period |
| --- | --- | --- | --- |
|  | Platelets | 24 | N.A. |
|  | Concurrent plasma | 12 | Total Plasma Loss (column I) - 14.4 liters (14,400 mL) if the donor weighs 175 lbs. or more and - 12.0 liters (12,000 mL) if the donor weighs 110- 174 lbs. These numbers exclude anticoagulant - A donor must not undergo more than 24 Platelet Pheresis collections in a 12 month rolling period. - The interval between each collection of Platelet Pheresis must be at least two days with no more than two procedures in a seven-day period. - Concurrent plasma must not be collected more than every 28 days. - A donor must not undergo more than 12 concurrent plasma collections in a 12 month rolling period. - Donors must not donate red blood cells for 8 weeks/56 days after a single RBC apheresis or 16 weeks/112 days after a double RBC apheresis donation |
|  | Red Cell | 3 double 7 whole blood | Total RBC Loss (column D) 1733 mL, based on the following calculation: (1733 = 45% x 550 mL (488 mL draw + 62 mL sample pouch + tubing. Volume of a maximum Unit of Blood (550mL = 500 + 10%)] |
|  | Continued on next page |  |  |

| ## Procedure | Perform the following steps to complete the Volume Loss Worksheet (VLW) |  |  |
| --- | --- | --- | --- |
|  | ## Step | Action |  |
| Cont’d | 11 | Product Yield | Transcribe the Mother bag Product Yield from Quality Control of Platelets, Pheresis form |
|  | 12 | Comments | Any comments on donation or condition of donor. |
|  | 13 | Initial, Staff | The staff member completing the form - If the donor requires a SU deferral, enter it at this time and make a comment in the comments section, SU entered. Enter the deferral code (SU) on the DHQ and the effective date and eligibility date, |
|  | 14 | Reviewed by/Date | Licensed staff prior to the next donation. - Eligibility is correct based on calculations below - Red Cell Loss and Plasma Loss calculations are accurate. - Number of donations is correct SU deferral added to SafeTrace as applicable-(if the eligibility date was incorrectly calculated or if an SU deferral was added to the last donation) |

| Controlled Documents | Attachment A: | Trima Accel Automated Blood Collection System Volume Loss Worksheet |
| --- | --- | --- |
|  | Attachment B: | Trima Accel Automated Blood Component System Procedure Runsheet |

| Authors: | Cecilia Militante, CLS | Carmen Flores, CLS |
| --- | --- | --- |
|  | Nora Manuel RN, MN | Angela Varela, MT(ASCP) |
|  | Jocelyn Guillermo, RN | Howard Ammerman, MT(ASCP) |
|  | Lynn Douglas | Lourdes Ramos, RN |
|  | Virginia Vengelen-Tyler, MBA, MT(ASCP)SBB |  |

| Uncontrolled Documents | 1. Guidance for Industry and FDA Staff: FDA: Non-binding recommendations, Dec. 2007 2. Residual rbc and plasma loss retained in TRIMA tubing set (for rinseback and no rinseback) and the extracorporeal blood loss is found in the TRIMA operating manual and is specific to type of TRIMA kit used. (LRS Platelet/ Plasma/ RBC or RBC/Plasma) 3. Standards for Blood Banks and Transfusions Services, 28th ed., Carson, TH, ed. American Association of Blood Banks, 2012. |
| --- | --- |

| Distribution | All SCPMG Blood Donor Centers |
| --- | --- |
|  | Continued on next page |

Reviewed and approved by:

| Electronically Signed |  | February 13. 2008 |
| --- | --- | --- |
| Virginia Vengelen-Tyler, MBA, MT,ASCP(SBB), CQA(ASQ) Regional Blood Bank Compliance Officer |  | Date |
| Signature Collected Electronically |  | ## February 6, 2008 |
| Gary Gochman, MD, Medical Director –Orchard Donor Center |  | ## Date |
| Signature Collected Electronically |  | ## August 6, 2009 |
| Joseph Thompson, MD. Medical Director –Los Angeles Donor Center |  | ## Date |
| Signature Collected Electronically |  | February 13, 2008 |
| Jeffery D. Shiffer, MD. Medical Director –Woodland Hills Donor Center |  | ## Date |
| Signature Collected Electronically |  | ## December 17, 2010 |
| Adriana A. Bedoya, M.D., FCAP, FASCP Medical Director- San Diego Donor Center |  | ## Date |
|  | ## Continued on next page |  |

DOCUMENT HISTORY PAGE

Effective Date: February 13, 2008

| Change type: new, major, minor etc. | Changes Made to Document – Describe | Signature responsible person/Date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ Date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| New |  |  |  |  |  |
| Minor | - Added detail – listed requirements - Defined how to use the run sheet and Plasma loss - Added detail on rinseback - Gave more specific examples on how to calculate | Ginny Tyler 08/07/09 | N.A. | N.A. |  |
| Major | - Added Trima - Added apheresis double RBC - Updated Attachment B | Ginny Tyler 12/28/10 | All approved by 12/17/10 | N.A. |  |
| Minor | Added the FDA requirement to defer for 16 weeks if RBC loss is equal to or greater than 300 mL. | Ginny Tyler 628/11 | N.A. | N.A. |  |

IMP = Implemented

| MasterControl History of Change: |  |  |
| --- | --- | --- |
| Change type: new, major, minor etc. | Version # | Description of Change |
|  | 5 | - Added Hemoglobin - Changed UDHQ to DHQ - Corrected calculations to allow for 62 mL collection in the pouch. - Included the reviews of the donations and documentation on the the “Trima Accel Automated Blood Component Collection System: Volume Loss Worksheet - Add a column for the review on the Worksheet |