#### BECKMAN COULTER AU680 SAMPLE PROCESSING

| Purpose | The Beckman Coulter AU680 chemistry analyzer measures analytes in samples, in combination with appropriate reagents, calibrators, quality control material and other accessories. ____________________________________________________________ |
| --- | --- |
| Scope Policy | The Beckman Coulter Au680 system is designed to analyze serum or plasma, urine, CSF, peritoneal and pleural fluid. To analyze samples, the appropriate reagents, reference materials, and QC samples are required. This system is designed to discard the concentrated waste Liquids (mixture) and washing waste liquids (washing water and detergents) separately. ____________________________________________________________ Quality control must be processed and be acceptable every 8 hours before sample results can be released. |

Specimen         Specimen for testing are collected following the container ID required for

Collection         for each specific analytes.

____________________________________________________________

Specimen         If it is necessary to store the specimen, store under refrigeration at 2 to 8°C.

Storage

- Continued on next page

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

| Specimen rejection | - Follow policy on Criteria for Unacceptable specimen (LGM 2030) |
| --- | --- |

| ##### Reagents and/or Media | Description | Vendor | Storage |
| --- | --- | --- | --- |
|  | Chemistry Cartridge reagents | Beckman Coulter | Refrigerated |
|  | ISE Buffer Solution | Beckman Coulter | Room temperature |
|  | ISE Mid Standard Solution | Beckman Coulter | Room temperature |
|  | ISE Reference Solution | Beckman Coulter | Room temperature |
|  | Wash Solution | Beckman Coulter | Room temperature |

| ##### Materials and supplies | - Disposable pipette tips - Sample cups, Hitachi cups (No. 716-0425) - Micro-sample cup, Hitachi cups (No. 707-0313) - Nested cups, Hitachi cups (No. 716-0415) - Sample Racks |
| --- | --- |

- Continued on next page

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

| ##### Safety or Special Safety Precautions | All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety and the safety of others and adhering to all departmental and medical center safety policies and procedures. - For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). - For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. - For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing. - For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. - For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste. - For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. - For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. - For proper storage and disposal of chemical hazardous waste, see LGM 8012. . |
| --- | --- |

| ##### Definitions | The Beckman Coulter AU680 is an automated chemistry analyzer that follows the analytic principle of Spectrum by a diffraction grating. |
| --- | --- |

- Continued on next page

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

| ##### Procedure: | - Follow the steps required for daily Start-up. |  |  |
| --- | --- | --- | --- |
|  | Step | Action |  |
|  | 1 | From Home screen select Start Condition jump button. |  |
|  | 2 | Select Edit F1. |  |
|  | 3 | Select New Index button. |  |
|  | 4 | Select Confirm F1. |  |
|  | 5 | Select OK to create a new index. (This procedure is to be performed daily right after midnight) |  |
|  | 6 | Confirm the analyzer status by selecting Analyzer Status from the home window. |  |
|  | 7 | Investigate any yellow or red display colors. |  |
|  | 8 | Perform all Daily Maintenance required. |  |
|  | 9 | Do a reagent check by selecting Reagent Management from the Home screen. |  |
|  | 10 | Select Reagent Check F5> Check All Positions > Start |  |
|  | 11 | View reagent information - From the Main Tab, view shots/volume available - From the Details Tab, view reagent stability and verify fixed reagents are in correct positions. - Repeat for each sample type. |  |
|  | 12 | Load new reagents if needed, then repeat step 10. |  |
|  | 13 | View reagent information to verify reagents have adequate stability and volume. |  |
|  | 14 | Do an ISE Start Up (as required) |  |
|  | 15 | Verify the ISE reagents are in date (90 day open bottle stability) and have adequate volume. |  |
|  | 16 | Load reagents as needed and prime. |  |

Note: Samples loaded on the Direct Tract Sampling will automatically be processed

by the analyzer. Any repeats or reruns will be manually performed by CLS.

- Continued on next page

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

| Procedure | Loading samples without Direct Tract Sampling Step; Action. 1; Before starting analysis, place samples in the correct racks. - White: Routine patient analysis and automatic repeat runs. - Yellow: Calibration analysis. - Green: QC analysis - Orange: Manual repeat run analysis - Red: Emergency analysis - Blue: Reagent blank analysis. 2; Place sample with the barcode label aligned in the open slot in the rack. Make sure all samples are decapped.. 3; Open the sample protective cover on the rack supply unit.. 4; Set the prepared racks on the racks supply unit with the rack barcode number facing the back of the instrument.. 5; Note: Place the rack to the right of the rack identification sensor (white board). 6; Select Start from the keyboard or software screen.. 7; When the screen labelled Start appears, verify that no red messages are on the Error List. If none, then select Start.. 8; When done, remove the sample racks from the rack collection area and cap samples.. |  |
| --- | --- | --- |
|  |  | - Manual Programming for Samples on Racks - Step; - Action. - 1; - Select Home. - 2; - Select Rack Requisition Sample jump button. - 3; - Select Sample button. - 4; - Select Test Requisition tab. - 5; - Select Switch button to select “Sample Kind”. - 6; - Select sample type from the “type” drop-down list. - 7; - Select Start Entry F1.. - 8; - Manually requisition sample, enter “Sample ID” (barcode number). - 9; - Select the test(s) to run on the sample. Selected tests are highlighted in blue. |

- Continued on next page

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

Procedure            Continued: Manual Programming for Samples on Racks

| 10 | If a manual dilution was made, select Sample dilution F7 and enter the dilution rate. |
| --- | --- |
| 11 | Select Demographics tab to enter required patient demographic information. |
| 12 | Select Entry F1 to save requisition. |
| 13 | Select Exit F2. |
| 14 | Select Pending List F4 to view a list of samples requisitioned |
| 15 | Select Close to close the pending List |
| 16 | Place the samples with barcodes in appropriate rack(s) |
| 17 | Select Start from main button bar |
| 18 | Review errors on the Error List in the Start Window and perform any corrective actions, if necessary. |
| 19 | Select Start from the Start Window |

Manual Programming for STAT Samples:

| 1 | Select Home |
| --- | --- |
| 2 | Select Stat Status jump button |
| 3 | Select Sample button |
| 4 | Select Test Requisition tab |
| 5 | Select sample type from the “Type” drop-down list |
| 6 | Select Start Entry F1 |
| 7 | Manually requisition sample: Enter “Sample ID” (barcode number) |
| 8 | Select the test(s) to run on the sample. Selected tests are highlighted in blue |
| 9 | If a manual dilution was made, select Sample dilution F7 and enter the dilution rate. |
| 10 | Select Demographics tab to enter required patient demographic information. |
| 11 | Select Entry F1 to save requisition. |
| 12 | Select Exit F2. |
| 13 | Select Pending List F4 to view a list of samples requisitioned |
| 14 | Select Close to close the pending List |
| 15 | Select STAT Status button |

- Continued on next page

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

Procedure            Continued: Manual Programming for STAT Samples:

| 16 | Check the “Kind” and “Type” columns for position assignment on the STAT table. |
| --- | --- |
| 17 | If the STAT table is busy (amber light flashing), Select STAT Pause F2, and then select OK. Wait for the amber light to stop flashing. |
| 18 | Open the small cover of the STAT table and load samples in appropriate positions with barcode facing out. Press the green TABLE ROTATION/DIAG button to rotate the table. |
| 19 | Close the small cover then select Start F1 |
| 20 | Review errors on the Error List in the STAT Start Window and perform any corrective actions, if necessary |
| 21 | Select Start from the STAT Start Window. |

Perform Add On and Rerun Tests on Rack:

| 1 | Select Home |
| --- | --- |
| 2 | Select Rack Requisition Sample jump button |
| 3 | Select Sample button |
| 4 | Select Test Requisition tab |
| 5 | Requisition on add on or rerun test: Select Add on F5 |
| 6 | Verify the correct “Sample Kind” rack in which the sample was initially processed is displayed. Select the Switch button if you need to change the sample kind. |
| 7 | Select sample type from the “Type” drop-down list |
| 8 | Enter the sample number in both of the “Sample No.” fields. Note: Enter a range of sample numbers if an add on or rerun is required on multiple samples for the same tests. |
| 9 | Select the Select Tests to be Repeated option |
| 10 | Select the test(s) to add on or rerun |
| 11 | Select OK |
| 12 | Select Pending List F4 to view a list of samples requisitioned. |
| 13 | Optional: Select the sample from the Pending List and select Go to view the specific sample requisition. Tests with an asterisk (*) are pending processing. |
| 14 | Select Close to close the Pending List |
| 15 | Place the samples with the barcodes in appropriate rack(s) and place them on the rack supply unit |
| 16 | Select Start from the main button bar |
| 17 | Select Start from the Start Window |

- Continued on next page

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

Procedure               Perform Add On and Rerun Tests on STAT Table:

Note: This procedure applies only for STAT samples that require add on or

Rerun test to be processed on the STAT table using the same sample ID and

Index as the original sample.

| 1 | Select Home |
| --- | --- |
| 2 | Select STAT Status jump button |
| 3 | Select Sample button |
| 4 | Select Test Requisition tab |
| 5 | Requisition on add on or rerun test: Select Add on F5 |
| 6 | Select sample type from the “Type” drop-down list |
| 7 | Enter the sample number in both of the “Sample No.” fields. Note: Enter a range of sample numbers if an add on or rerun is required on multiple samples for the same tests. |
| 8 | Select the Select Tests to be Repeated option |
| 9 | Select the test(s) to add on or rerun |
| 10 | Select OK |
| 11 | Select Pending List F4 to view a list of samples requisitioned. |
| 12 | Optional: Select the sample from the Pending List and select Go to view the specific sample requisition. Tests with an asterisk (*) are pending processing. |
| 13 | Select Close to close the Pending List |
| 14 | Select STAT Status button |
| 15 | Check the “Kind” and “Type” columns for position assignments on the STAT Table |
| 16 | If the STAT table is busy (amber light flashing) Select STAT Pause F2, then select OK |
| 17 | Wait for the amber light to stop flashing |
| 18 | Open the small cover of the STAT table and load samples in appropriate positions with barcodes facing out. Press the green TABLE ROTATION/DIAG button to rotate the table then close the small cover. |
| 19 | Select STAT Start F1 |
| 20 | Select Start from the STAT Start Window. |

- Continued on next page

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

#### Procedure             Recall Patient Results

| 1 | Select Home |
| --- | --- |
| 2 | Select Sample Manager jump button. Note: Select OK if a “Data Not Found” message appears |
| 3 | Select Main Tab |
| 4 | Select Data Search F3 |
| 5 | Select start date of search in the “Start Index” drop-down list |
| 6 | Select end date of search in the “End Index” drop-down list |
| 7 | Select Search the designated sample option |
| 8 | Select the check box for the desired sample kind/sample type from the “Sample Kind” column. Note: You may need to select the drop down arrow to view more sample kind options. |
| 9 | Select Complete Match from the “Search Sample ID” drop-down list |
| 10 | Enter Sample ID in “Search Sample ID” field |
| 11 | Select OK |
| 12 | To view report, select Sample tab |
| 13 | To print report, select Print F8 the select OK. |

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

Controlled            The controlled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
| LGM 8000 | Safety Practices |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |
| LGM 8006 | Handling of Regular and Infectious Waste |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand-washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |
| LGM 294 | Beckman Coulter AU 680 Quality Control |

Non

Controlled            The non-controlled documents are as follows:

Documents

| Document No. | Document Name |
| --- | --- |
|  | AU680 User’s Guide |

Author                    Aida R. Legaspi. CLS

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Aida Legaspi, CLS Area Laboratory Manager – Los Angeles Medical Center |  |
| Julie Toti, CLS, MS Laboratory Director – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

#### BECKMAN COULTER AU680 SAMPLE PROCESSING

___________________________________________________________

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |