#### URINALYSIS AUTORELEASE PROCESS

| Process | With autorelease LMS of urinalysis is not required. All urinalysis samples not requiring microscopic analysis are autorleased from ICHEM Velocity. Samples requiring microscopic analysis are released from Iris IQ200. This document describes the auto-release process of urinalysis results. |
| --- | --- |

| ##### Work place safety | All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, and the safety of others and adhering to all departmental and medical center safety policies and procedures. - For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). - For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. - For proper hand-washing, see LGM 8010, specifically, not limited to, proper hand-washing. - For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. - For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and bio hazardous waste. - For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. - For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. - For proper storage and disposal of chemical hazardous waste, see LGM 8012. |
| --- | --- |

Continued on next page

URINALYSIS AUTORELEASE PROCESS, Continued

| Procedure | Following is the process for urinalysis auto-release |
| --- | --- |

| Step | Action |
| --- | --- |
| 1 | Verify that daily QC is done and acceptable. (note: If QC is not acceptable you must resolve QC and have acceptable QC results prior to running any patient samples). If QC is acceptable you may then auto-release urinalysis results as follows. |
| 2 | Load Samples on IRIS Velocity |
| 3 | Racks automatically move from IRIS Velocity to IQ200 Sprint. All Urine Chemistry results will be autoverified except for Glucose with result of >500mg/dL and patient is <18 years old. This result will be a critical value and is intended to aid in the identification of new type 1 diabetes mellitus patients presenting in diabetic ketoacidosis. |
| 4 | Perform critical call procedure in Cerner for patients with glucose of >500mg/dL and is <18 years old. |
| 5 | Samples that are positive Urine Chemistries such as blood, leukocyte esterase, nitrite or protein >trace will reflex a UA microscopic and will be analyzed by IRIS IQ200 for microscopic findings. |
| 6 | All microscopic verified from IRIS IQ200 will autoverify in cerner. |

Continued on next page

URINALYSIS AUTORELEASE PROCESS, Continued

| ##### Controlled documents | See below |
| --- | --- |

| Documents | Title |
| --- | --- |
| LGM 8000 | Safety Practices |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |
| LGM 8006 | Handling of Regular and Infectious Waste |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand-washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |
| LIMM209.2 | Results entry specimen results processing. |
| LUM252 | ICHEM Velocity Sample Processing |
| LUM242 | Specimen Processing in IQ200 |

| ##### Author | Alvin Castillo, CLS |
| --- | --- |