#### Beckman Coulter Unicel® DxH 800 System®

XB Analysis

PURPOSE To give the instructions for monitoring XB on Beckman Coulter Unicel DxH 800 System.

| ##### PRINCIPLE | The XB Analysis is a quality control method that monitors instrument performance (calibration) by tracking the MCV, MCH, and MCHC parameters of all patient samples. The XB will be monitored on each shift for out of control, shift, and trends. |
| --- | --- |

Workplace	    All laboratory employees are expected to maintain a safe working environment

Safety		    and an injury-free workplace.  Laboratory employees are responsible for their

own safety, and the safety of others and adhering to all departmental and

medical center safety policies and procedures.

- For standard precautions and safety practices in the laboratory;  see LGM

8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids.

- For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing.

- For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves.

- For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and bio hazardous waste.

- For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas

Continued on next page

#### Beckman Coulter Unicel® DxH 800 System®

XB ANALYSIS

Continued,

- For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- For proper storage and disposal of chemical hazardous waste, see LGM 8012.

| ##### DEFINITION | Out of Control Batch – the XB analysis is out of control when the batch mean is not within the established limits of the target mean. What does an “Out of Control” Batch mean? - An instrument problem - A reagent problem - A sample handling problem - A calibration problem - A change in the patient population. |
| --- | --- |

ACTION

REEQUIREMENTS

- Out of Control Batch – the technologist must evaluate the following criteria to determine if the instrument can be operated.

- Shift Detected – if a shift is detected but results are within limits, continue to operate, however:

| Step | Action |
| --- | --- |
| 1. | Notify Manager/Designee. |
| 2. | Monitor XB results during the shift to ensure that they remain within limits. |

Continued on next page

#### Beckman Coulter Unicel® DxH 800 System®

XB ANALYSIS

Continued,

| Step | Action |
| --- | --- |
| 3. | If problem is detected, correct/troubleshoot and continue to monitor. |
| 4. | Document under Comment Icon on the Local Navigation Bar. |

- Trend Detected – if a trend is detected and results are within limits continue to operate, however:

| Step | Action |
| --- | --- |
| 1 | Notifiy Manager/Designee. |
| 2. | Monitor the XB results during the shift to ensure that it remains within limts |
| 3 | If a problem is detected, correct/troubleshoot and continue to monitor. |
| 4. | Document under Comment Icon on the Local Navigation Bar. |

_________________________________________________

| ##### TROUBLE SHOOTING | - One batch is out of limits but no shift or trend is noted. Run Controls. If controls are in limits, continue to operate but monitor each XB batch. If XB continues to be out three or more batches, take remedial action or use back up procedures. - One batch is out of limits and previous batches have demonstrated shift or trends which indicate that the out of limits is a result of an ongoing problem. Take remedial action use back up procedures. - More than one batch is out of limits. Take remedial action or use back up procedures. |
| --- | --- |

_______________________________________________

#### Beckman Coulter Unicel® DxH 800 System®

XB ANALYSIS

Continued,

| ##### REVIEWING XB/XM RESULTS | Select XB/XM on the Alert Status Icon to display XB Batch Means. Click the XB icon twice and XM Batch Means will be displayed. The workstation determines if the results should be included in a batch by considering several factors. - RBC value is less than 1.0 Mil. - Flow cell clog or partial aspiration - Non numeric value for MCV, MCH, and MCHC - Non numeric value for MCV, MCH, and MCHC and at least 25% of the samples have been excluded for MCV, or MCH, and or MCHC, or manually excluded by the operator. |
| --- | --- |

_________________________________________________________________

Controlled	The controlled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
| LGM 2022 | Quality Manual |
| LGM 8000 | Safety Practices |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |
| LGM 8006 | Handling of Regular and Infectious Waste |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |

#### Beckman Coulter Unicel® DxH 800 System®

XB ANALYSIS

Continued,

NON			The Non Controlled Documents are as follows:

CONTROLLED

DOCUMENTS

| Document No. | Name of Documents |
| --- | --- |
|  | Beckman Unicel DxH Coulter Operator’s Manual |

__________________________________________________________________

AUTHOR		Marianne E. Muyrong, CLS

#### Beckman Coulter Unicel® DxH 800 System®

XB ANALYSIS

Continued,

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Julie Toti, CLS, MS Manager, Clinical Laboratory – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

Continued on next page

#### Beckman Coulter Unicel® DxH 800 System®

XB ANALYSIS

Continued,

## HISTORY PAGE

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Laboratory Director review/date | Laboratory Manager review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New | . |  |  |  |  |