#### Beckman Coulter UniCel DxH 800  &  DxH Slidemaker Stainer

#### Daily Checks and Shutdown

| ##### Purpose Policy | To provide instructions on how to perform daily, weekly and as needed maintenance for Beckman Coulter Unicel DxH 800 and DxH Slidemaker Stainer. . _____________________________________________________________ Shutdown procedure and Daily Check diagnostic tests are automatically performed by the instrument at the configured time every day and must be documented on the Daily Maintenance Log with the CLS initial who performed the procedure. Shutdown and Daily Check may also be performed when needed for troubleshooting purposes or when a Shutdown or Daily check is unsuccessfully performed. |
| --- | --- |
| ##### Scope | _____________________________________________________________ Maintenance includes but not limited to: - Performing Shutdown and Daily Checks - Replacing Reagents - Cleaning Procedure |

| ##### Workplace Safety | All laboratory employees are expected to maintain a safe working environment Safety and an injury-free workplace. Laboratory employees are responsible for their own safety, and the safety of others and adhering to all departmental and medical center safety policies and procedures. • For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). • For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. |
| --- | --- |

#### Beckman Coulter®UniCel® DxH 800  & Slidemaker Stainer

#### Daily Checks and Shutdown

Continued,

- For proper hand-washing, see LGM 8010, specifically, not limited to, proper hand-washing.

- For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves.

- For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste.

- For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas.

- For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- For proper storage and disposal of chemical hazardous waste, see LGM 8012.All laboratory employees are expected to maintain a safe working environment and an injury-free workplace.  Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures.

__________________________________________________________________

PROVISIONS/PROCEDURE

- LOG OFF/		In order to access the System Manager, log on/log off is necessary:

- LOG ON

| Step | Action |
| --- | --- |
| 1. | Select Log off icon. |
| 2. | On the Log off screen, select the log off button |
| 3. | On the Logon screen, type the Operator ID (case sensitive) |
| 4. | Type in Password (case sensitive) |
| 5. | Select Log on button or Press Enter. |
| 6. | Click Start Icon to place DxH System online for processing ->Stop Icon to place DxH System offline (inactive state), ->Alarm Icon to shut the audible alarm off. |

- __________________________________________________________________

- DAILY 		The DxH System is programmed to automatically perform a 30 minute

- SHUTDOWN	shutdown every 24 hours to ensure optimal cleaning of the instrument.

- The schedule for the daily shutdown is as follows:

- DxH 800 #1  and SMS #1 – every 3:00am

- DxH 800#2 and SMS #2 – every 1:00am

- DxH 800#3 – every 2:00am

#### Beckman Coulter®UniCel® DxH 800  & Slidemaker Stainer

#### Daily Checks and Shutdown

Continued,

__________________________________________________________________

In any event that the Daily Shutdown was unsuccessfully performed or when a technical support requests it to be done during troubleshooting,  a shutdown may be manually performed using the following steps:

| Step | Action |
| --- | --- |
| 1. | Click Daily Check Icon |
| 2. | Select Shutdown (Dialog box displays). |
| 3. | Select Perform Daily Check After Shutdown. |
| 4. | Enter a specific time (default is 30 minutes, may be requested to a maximum of 24 hours). |
| 5. | Select OK to begin shutdown. |

NOTE:  Once shutdown is in process, no other cycles or System Manager Actions can be performed until Shutdown cycle has completed.  If shutdown needs to be cancelled for any reason:

- Click Cancel

- Perform Daily Check.

ADDITIONAL TASKS - Prime the SPM to remove cleaner in case an error occurred during a Shutdown cycle. Select the Prime button at the bottom of the Daily Checks screen.  Document on Daily Maintenance Log with CLS initial, that Daily Shutdown was successfully performed.

__________________________________________________________________

| - DAILY CHECKS | After the daily shutdown, DxH System is programmed automatically to perform Daily Checks (refer to the DxH System schedule of Daily Shutdown above). If the Daily Checks alert status icon is in neutral color and review button is disabled, the Daily Checks pass. Daily Checks may also be manually requested in the instrument during troubleshooting/when requested by a technical support or when changing reagents using the the following steps: |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Click Daily Check Icon to display Daily Check Summary. |
| 2. | Select Daily Checks button at the local navigation bar (a diaolog box will display). |

#### Beckman Coulter®UniCel® DxH 800  & Slidemaker Stainer

#### Daily Checks and Shutdown

Continued,

|  | Step; Action. 3.; Select OK to run Daily Checks (Cycle will run approx. 4.5 minutes for DxH 800 and 3minutes for DxH SMS).. NOTE; If Daily Check Alert status icon is in red, then something failed and needs review. When this happens, you need to acknowledge the value to return the status color to neutral to be able to run the instrument. - Look at each Daily checks tab, beginning with the Summary tab. - If any background result fails, run Daily checks again. - Verify background results when the Daily checks cycle finishes. - Add a comment (refer to instructions below).. |
| --- | --- |

| ADDING A Step | Action |
| --- | --- |
| 1. | Select View Log from the Local Navigation Bar. |
| 3. | Select one data at a time to review |
| 4. | Selected data will be highlighted in Blue. |
| 6. | Select Comment button, then comment/corrective action taken. |
| 7. | Enter comment/corrective action taken. |
| 8 | Select OK to save comment. Once saved a “yellow” icon will appear. |

COMMENT

#### Beckman Coulter UniCel DxH 800  &  DxH Slidemaker Stainer

#### Daily Checks and Shutdown

Continued,

ADDITIONAL

TASKS	Prime the SPM. Prime provides a method to remove cleaner in case an error occurred during a Shutdown cycle. Select the Prime button at the bottom of the Daily Checks screen to prime the SPM.

| ##### Controlled Documents | The table below lists the controlled documents Document Number; Document Name. LGM 8000; Safety Practices. LGM 8004; Infection Control. LGM 8005; Universal Body Substance Precaution. LGM 8006; Handling of Regular and Infectious Waste. LGM 8007; Cleaning Work Areas. LGM 8010; Hand-washing Policy. LGM 8012; Storage and Disposal of Chemical Hazardous Waste. LGM 2022; Quality Control Policy. |
| --- | --- |

| ##### Uncontrolled Documents | The table below lists the uncontrolled documents Document Number; Document Name. Instructions for Use UniCel® DxH 800 Coulter® Cellular Analysis System. |
| --- | --- |

| ##### Author | Marianne Muyrong, CLS |
| --- | --- |

Continued on next page

#### Beckman Coulter UniCel DxH 800  &  DxH Slidemaker Stainer

#### Daily Checks and Shutdown

Continued,

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Julie Toti, CLS, MS Manager, Clinical Laboratory – Los Angeles Medical Center |  |
| Joseph C. Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

Continued on next page

#### Beckman Coulter UniCel DxH 800  &  DxH Slidemaker Stainer

#### Daily Checks and Shutdown

Continued,

#### HISTORY PAGE

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Laboratory Director review/date | Laboratory Manager review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  |  |  |  |  |