#### Beckman Coulter UniCel® DxH 800 System

#### Instrument Function Checks

| ##### POLICY | If a testing instrument does not meet function checks or performance testing requirement, this same instrument MUST NOT BE USED TO REPORT PATIENT RESULTS. Corrective action must be performed and documented to remedy the problem before analyzing patient samples. |  |
| --- | --- | --- |

| ##### Workplace Safety | All laboratory employees are expected to maintain a safe working environment Safety and an injury-free workplace. Laboratory employees are responsible for their own safety, and the safety of others and adhering to all departmental and medical center safety policies and procedures. • For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). • For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. • For proper hand-washing, see LGM 8010, specifically, not limited to, proper hand-washing. • For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. • For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste. • For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. • For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. • For proper storage and disposal of chemical hazardous waste, see LGM 8012.All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures. |
| --- | --- |

Continued on next page

#### Beckman Coulter UniCel® DxH 800 System

#### Instrument Function Checks

#### Continuation,

- PROCEDURE

| STEPS | ACTION |
| --- | --- |
| 1 | On the Alert Status Icons; click the History Log. |
| 2 | Select the high-lighted message. |
| 3 | Select “DETAILS” and click on ‘TROUBLESHOOT” tab. |
| 4 | Help screen will pop-up and follow suggested corrective actions. |
| 5 | Once corrected; perform a daily checks, run Latron CP-X and all 6 C and Retic-X controls. Document actions at the control comment section and at the Hematology Maintenance Log. All parameters must be within the limits before a specimen is performed. For further monitoring, notify next shift CLS and Manager. |
| 6 | If problem is not corrected, go back to troubleshooting and follow the next suggested corrective action or call Technical Support for help. |
| 7 | Once the problem is corrected, repeat other patient results from same testing sequence to ensure accuracy of result and/or correct if necessary following Cerner GenLab correction of results. |

| ##### Controlled Documents | The table below lists the controlled documents Document Number; Document Name. LGM 8000; Safety Practices. LGM 8004; Infection Control. LGM 8005; Universal Body Substance Precaution. LGM 8006; Handling of Regular and Infectious Waste. LGM 8007; Cleaning Work Areas. LGM 8010; Hand-washing Policy. LGM 8012; Storage and Disposal of Chemical Hazardous Waste. LGM 2022; Quality Control Policy. |
| --- | --- |

| ##### Author | Marianne Muyrong, CLS |
| --- | --- |

| ##### Uncontrolled Documents | The table below lists the uncontrolled documents |
| --- | --- |

| Document Number | Document Name |
| --- | --- |
|  | Instructions for Use UniCel® DxH 800 Coulter® Cellular Analysis System |

#### Beckman Coulter UniCel® DxH 800 System

#### Instrument Function Checks

#### Continuation,

__________________________________________________________________

#### Beckman Coulter UniCel® DxH 800 System

#### Instrument Function Checks

#### Continuation,

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Julie Toti, CLS, MS Manager, Clinical Laboratory – Los Angeles Medical Center |  |
| Joseph C. Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

Continued on next page

#### Beckman Coulter UniCel® DxH 800 System

#### Instrument Function Checks

#### Continuation,

HISTORY PAGE

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Laboratory Director review/date | Laboratory Manager review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  |  |  |  |  |