#### Beckman Coulter Unicel® DxH Slidemaker Stainer

#### Cleaning Procedure

| ##### Purpose | The purpose is to give instructions for performing the Cleaning Procedure for the Unicel DxH Slidemaker Stainer Coulter. |
| --- | --- |

| ##### Policy | Unicel DxH Slidemaker Stainer Cleaning Procedure must be performed |  |
| --- | --- | --- |

| ##### Workplace Safety | All laboratory employees are expected to maintain a safe working environment Safety and an injury-free workplace. Laboratory employees are responsible for their own safety, and the safety of others and adhering to all departmental and medical center safety policies and procedures. • For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). • For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. • For proper hand-washing, see LGM 8010, specifically, not limited to, proper hand-washing. • For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. • For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste. • For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. • For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. • For proper storage and disposal of chemical hazardous waste, see LGM 8012.All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures. |
| --- | --- |

Continued on next page

#### Beckman Coulter Unicel® DxH Slidemaker Stainer

#### Cleaning Procedure

Continuation,

- PROVISIONS/PROCEDURE

|  | Follow the steps below to complete the Cleaning Procedure |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Drain all baths |
| 2. | Transfer pickup tubes of supplies 2 and 4 to the Methanol cleaning bottles. Wrap pickup tubes assembly in clean gauze to prevent dripping. |
| 3. | Individually fill baths (2 and 3) with Methanol. |
| 4. | Allow the Methanol to sit in the baths for at least 15 minutes. |
| 5. | Drain baths (2 and 3) individually or you may select Drain all Baths. |
| 6. | Remove pick up tubes assembly from the Methanol cleaning bottles and turn them upside down. Wrap pick up assembly in clean gauze to prevent from dripping. |
| 7. | While holding pick-up tubes individually fill baths (2 and 3) with air. Allow air to flow through pick –up tubes for less than 2 minutes and press Cancel. (This will prevent the low supply warning). |
| 8. | Transfer pick up tubes assembly back into supplies ( 2 and 4) |
| 9. | Select the Finish icon from the Local Navigation Bar. |
| 10. | Clean or Replace baths, especially baths (2 and 3). |
| 11 | Clean the probes by wiping them with Methanol. |
| 12. | Place DxH Slidemaker Stainer Online to continue working. |
| 13. | Document on DxH Slidemaker Stainer Maintenance Log the date cleaning was performed and CLS initial. |

Continued on next page

#### Beckman Coulter Unicel® DxH Slidemaker Stainer

#### Cleaning Procedure

Continuation,

| - MATERIALS - NEEDED | Gloves Gauze Pads Methanol Lab Coat |
| --- | --- |

| ##### Controlled Documents | The table below lists the controlled documents Document Number; Document Name. LGM 8000; Safety Practices. LGM 8004; Infection Control. LGM 8005; Universal Body Substance Precaution. LGM 8006; Handling of Regular and Infectious Waste. LGM 8007; Cleaning Work Areas. LGM 8010; Hand-washing Policy. LGM 8012; Storage and Disposal of Chemical Hazardous Waste. LGM 2022; Quality Control Policy. |
| --- | --- |

| ##### Uncontrolled Documents | The table below lists the uncontrolled documents Document Number; Document Name. Instructions for Use UniCel® DxH 800 Coulter® Cellular Analysis System. |
| --- | --- |

| ##### Author | Marianne Muyrong, CLS |
| --- | --- |

Continued on next page

#### Beckman Coulter Unicel® DxH Slidemaker Stainer

#### Cleaning Procedure

Continuation,

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Julie Toti, CLS, MS Manager, Clinical Laboratory – Los Angeles Medical Center |  |
| Joseph C. Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

Continued on next page

#### Beckman Coulter Unicel® DxH Slidemaker Stainer

#### Cleaning Procedure

Continuation,

#### HISTORY PAGE

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Laboratory Director review/date | Laboratory Manager review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  |  |  |  |  |