IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

PURPOSE	The purpose is to give instructions and guidelines on performing maintenance for the Iris Alifax Test 1 ESR Analyzer.

______________________________________________________________

##### POLICY	Maintenance must be performed in accordance to the guideline provided in this  procedure to ensure the analyzer’s optimum performance.

_____________________________________________________________

WORKPLACE	All laboratory employees are expected to maintain a safe working

SAFETY	environment and an injury-free workplace.  Laboratory employees are responsible for their own safety, and the safety of others and adhering to all departmental and medical center safety policies and procedures.

- For standard precautions and safety practices in the laboratory;  see LGM

8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids.

- For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing.

- For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves.

- For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and bio hazardous waste.

- For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- For proper storage and disposal of chemical hazardous waste, see LGM 8012.

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

Gloves, wash tubes

MATERIALS	0.5% Bleach

##### DI Water

Printer Paper

Empty Test tubes

Universal Card

Waste Tank Container

SWITCH ON

INSTRUMENT

| Step | Action |
| --- | --- |
| 1. | Verify that the front door is closed. |
| 2. | Press the rear side switch to turn on analyzer. |
| 3. | Wait at least 2 minutes before starting analysis to allow temperature to reach 37 C to get its stability |
| 4. | When activated, the instrument is going to prints and gives two acoustic beeps. And then proceed with initialization phase moving all mechanical arts. After which printer prints the following information: - The installed SW version. - Serial Number of the instrument - Date & Time - Test available inside the instrument to perform ESR analysis. |

_________________________________________________________________

TANK EMPTYING

OR REPLACEMENT

Warning! A SW function calculates the volume of blood wasted into the waste

tank after analysis. A counter is going to increase each processed sample and the

number of analysis that can be processes before activating the waste tank alarm is 			set to 1500

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

| Step | Action |
| --- | --- |
| 1. | Analyzer displays 0 = NOT EMPTY; 1 = EMPTY. |
| 2. | If Key 0 (Not Empty) = allows the instrument to go on counting for the further discarded tests. |
| 3. | If Key 1 (Empty) to reset the wated samples counter to 0. |
| 4. | The reset process can also be recalled manually by pressing Key 9 which is Tank Replace on Main Menu and then 9. |
| NOTE | For the disposal of the waste tank content, follow the standard safety procedures in use in the laboratory. |

__________________________________________________________________

MAIN MENU	To access the Operator Main Menu:

OPTION

| Step | Action |
| --- | --- |
| 1. | Press Key 8 (Operat. Menu) |
| 2. | Pres Key 1 – to print the list of the functions assigned to each key. |

________________________________________________________________________

CHANGING

DATE & TIME

| Step | Action |
| --- | --- |
| 1. | Press Key 4 on Main Menu |
| 2. | If date or time is correct, press ENTER. |
| 3. | If date or time needs to be changed, Press CLEAR. |
| 4. | Type the correct date. For the year, only the last two digits can be changed. |

_________________________________________________

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

| LOADING Step | Action |
| --- | --- |
| 1. | Open the cover of the printer and remove the residual part of paper. |
| 2 | Pull out roughly 15cm of paper from the roll, remove the part stuck on paper and be sure that the paper stream descends from the bottom side of the roll. |
| 3. | Place new roll into the printer lodgement with the paper border out from the printer. |
| 4. | Close the cover pressing it deeply until the two retaining clips are locked to the printer. |
| 5. | Press Paper Feed on the keypad or the icon on the printer to obtain more paper I it is necessary. |

PRINTER

PAPER

__________________________________________________________________

INCREASING		The instrument works with credits of test which can be loaded by means

TEST 			of Smart Card ALIFAX supply.  The amount of credits the customer can

AVAILABILITY	acquire is between 1000, 4000, 10000,20000.

| Step | Action |
| --- | --- |
| 1. | On Main Menu, press Key 0 |
| 2. | Press Key 1 |
| 3 | When “SM FST Area OK” is displayed, insert Card into the reader slot and press ENTER. |
| 4 | Printer will print-out a message with the number of tests available. |
| 5. | Printer will print-out a message with the number of tests available. |

____________________________________________________________

| HOW TO CHECK Step | Action |
| --- | --- |
| 1. | Press 2, and analyzer will ask to insert the card into the slot. |
| 2. | Insert Card and wait. |
| 3 | The display will show information according to the Smart Card status, either “Used Card” or “Not Used Card” |
| NOTE | To get Smart Card Log, Press 3 and history print-out cam be accessed. |

CARD

_____________________________________________

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

SMART		Malfunctions may be caused by:

CARD 			1. Smart card not inserted properly or inserted upside down.

ERRORS		2. The Card contact  plaque not set to the internal side of the instrument.

3. The reader contacts don’t allow the card to be read.

___________________________________________________________

WASHING		This procedure is designed to clean the complete capillary tubing to set it

PROCEDURE	free from blood or Latex residuals.  In order to  maintain the capillary cleaned and to increase its own working life, the operator should carry-out washings in different ways as described below:

- Washing Using 2 Test Tubes:

Used during the instrument working day when the instrument activates the warning to be washed by automatic request.

| Step | Action |
| --- | --- |
| 1. | Prepare 2 test tubes 3/4filled with distilled water. |
| 2. | Insert in position 1 and 2. |
| 3. | Press Key 2 (Washing) on Main Menu |
| 4. | Insert rack in the instrument. |
| 5. | Wait until Main Menu is displayed again, which point out the end of process. |

- Washing Using 3 Test Tubes:

Execute at the end of the working day to maintain the capillary cleaned during the night and therefore to increase its working life and to an easy way to remove residuals of blood particles from the needle and capillary tubing.

| Step | Action |
| --- | --- |
| 1. | Prepare 3 test tubes ¾ filled with distilled water. |
| 2. | Insert in the first three (3) positions. |
| 3. | Press Key 2 (Washing) on Main Menu. |
| 4. | Load the rack into the instrument. |
| 5. | Wait until “Test 1 off” message is displayed. |
| 6. | Press “Enter” if operator chooses to continue with running samples or operator can Switch the instrument off for the night. |

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

WSASHING	PROCEDURE

FOR MAINTENANCE

| Step | Action |
| --- | --- |
| 1. | Prepare tow (2) test tubes ¾ filled with distilled water. |
| 2. | Load in 1st and 2nd positions. |
| 3. | Press Key 2 (Washing) |
| 4. | Insert rack into the instrument and wait for the procedure to end. |
| 5. | Prepare one test tube ¾ filled with Sodium Hupochlorite (5% of dilution). |
| 6. | Load in 1st position. |
| 7. | Prepare one test tube ¾ filled with distilled water. |
| 8. | Load in 2nd position. |
| 9. | Press Key 2 (Washing). |
| 10. | Insert rack into the instrument and wait for the procedure to end. |
| 11. | In order to rinse the capillary, prepare two (2) test tubes ¾ filled with distilled water |
| 12. | Load in 1st and 2nd positions. |
| 13. | Press Key 2 (Washing), |
| 14. | Insert the rack into the instrument and wait for the procedure to end. |

NOTE: Washing procedure for maintenance is done when running daily controls.

At the end of each washing procedure, the software is going to report its value matched with suffix T100.  At every incorrect washing procedure, will generate Z error and a new washing procedure will be requested.  Document performance of  Washing Procedure for Maintenance on the Maintenance Log Sheet with date and CLS initial.

_________________________________________________

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

..

QUICK NEEDLE

REPLACEMENT

WARNINGS:  DO NOT touch the top of the syringe piston, because a light pressure could allow the needle to escape.  Moreover,  the needle tip could pierce the glove and the skin.  Operate with extreme caution and wear the appropriate PPE at all times.

| Step | Action |
| --- | --- |
| 1. | With instrument powered ON, open the front door to access the waste tank. |
| 2. | Rotate the metallic knob and open the metallic door to access the needle-syringe group. |
| 3. | Press Clear at the Main Menu, “Change Needle Enter Continue” message will be displayed. |
| 4. | Press Enter and the syringe group will move towards the front of the instrument. During this operation, the following message will be displayed “ Enter to End”. |
| 5. | After having unscrewed the retaining screw that locks the piston assembly to the support, grasp the piston with fingers and pull itowards you so that it will be inclined to 45 degrees. |
| 6. | Unscrew the piston cap and use the supplied tool if it is necessary. |
| 7. | Take the green tool applied over the plastic flab internal side and insert into the piston hole and unscrew to remove the needle. |
| 8. | Take the new kit (PN SI195077), with the needle inside, set over the plastic flap internal side and insert it to the piston hold and twist it. |
| 9. | After removing the green tool re-cap the piston |
| 10. | Push the piston assembly gently inside the instrument until the square magnets retains it vertically lock the piston assembly by twisting the retaining screw. |
| 11. | Lock the piston assembly by twisting the retaining screw. |
| 12. | Press Enter to reset the piston assembly to home position and close the doors. |
| 13 | Document with date replaced and CLS initial under As needed maintenance. |

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

Controlled	The controlled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
| LGM 2022 | Quality Manual |
| LGM 8000 | Safety Practices |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |
| LGM 8006 | Handling of Regular and Infectious Waste |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |

___________________________________________________________

##### NON			The Non Controlled Documents are as follows:

CONTROLLED

DOCUMENTS

| Document No. | Name of Documents |
| --- | --- |
|  | Iris Alifax Test 1 Analyzer Operator’s Manual |

___________________________________________________________

Author		Marianne Muyrong, CLS

___________________________________________________________

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Julie Toti, CLS, MS Manager, Clinical Laboratory – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

Continued on next page

IRIS ALIFAX TEST 1 ESR ANALYZER

MAINTENANCE

Continued..

.

## HISTORY PAGE

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Laboratory Director review/date | Laboratory Manager review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New | . |  |  |  |  |