

1

**LAMC LABORATORY**

|  |
| --- |
| ***Inter-Office Memorandum*** |
|  |  |  |  |
| **To:**  | ALL CLS  | **Date:** | 10/03/2014 |
| **From:**  | AIDA LEGASPI |  |  |
| **Subject:** | Aspirin and T Bili Body Fluid | **Location:** | LAMC |
|  |  | **Ext.:** |  |
| **CC:** |  |  |  |

Validation for Aspirin and Total Bili for Body fluid have been completed. Providers will be able to order the tests.

**Aspirin:**

* To be performed in our Verify Now instrument #1
* Use Aspirin kit
* Load sample as how you would load Plavix (P2Y12)
* Blood collection requirement and handling is the same as Plavix (P2Y12)
* Wait for 30 minutes after collection before running the sample
* Specimen stable up to 4 hours
* Result out in Cerner

**Total Bilirubin, Body Fluid**

* Sample should be kept on an amber container
* Load/program sample in Au680 manually
* Result in Cerner manually (do not use remisol)