IRIS ALIFAX TEST 1 ESR ANALYZER

QUALITY CONTROL

PURPOSE	The purpose is to give guidelines in running the Quality Control for the Iris Alifax Test 1

ESR Analyzer before a patient samples is performed.

##### POLICY	QC will be run once every 24 hours.  Running of QC will rotate every month and will be run at the beginning of the shift. Schedule of running Control is:

Day Shift: JANUARY, APRIL, JULY, OCTOBER

PM Shift: FEB, MAY, AUGUST, NOVEMBER

GY Shift: MARCH, JUNE, SEPTEMBER, DECEMBER

##### Control must be run after any scheduled or unscheduled preventive maintenance or repair, after calibration for calibration verification.  Controls are run as a patient, using the same reagents, and performed by the same personnel testing patient samples.  When quality control tolerance limits are exceeded, corrective action must be performed and recorded before analyzing patient samples.

Before use, controls must be stored between 4 to 25 C.  From the first piercing, keep refrigerated between 2 to 8 C and stable for 6 weeks.  Refrigerated control material should be returned to room temperature before running on the instrument.

Workplace	All laboratory employees are expected to maintain a safe working environment

Safety		and an injury-free workplace.  Laboratory employees are responsible for their

own safety, and the safety of others and adhering to all departmental and medical

center safety policies and procedures.

- For standard precautions and safety practices in the laboratory;  see LGM

8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids.

- For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing.

- For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves.

____________________________________________________________

Continued on next page

IRIS ALIFAX TEST 1 ESR ANALYZER

QUALITY CONTROL

Continued..

- For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and bio hazardous waste

- For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas

- For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- For proper storage and disposal of chemical hazardous waste, see LGM 8012.

MATERIALS	Latex Controls

##### Printer Paper

Gloves

3 Test Tubes ¾ filled  and 3 Test tubes filled with 3ml distilled water

1 Test Tube ¾ filled with 5% Sodium Hypochlorite

| BEFORE Step | Action |
| --- | --- |
| 1. | Execute first washing Procedure: - Press key 2 at Main Menu. - Using 2 test tubes, ¾ filled with distilled water, load tube in position 1 and 2 of the rack. |
| 2. | Execute second washing procedure: - Press key 2 at Main Menu. - Load 1 test tube filled with ¾ Sodium Hypochlorite in position 1 and 1 test tube filled with ¾ distilled water in position 2 of the rack. |

ANALYSIS

_________________________________________________________________

| Step | Action |
| --- | --- |
| 1. | Verify that the analyzer is on Main Menu. |

PROCEDURE

____________________________________________________________

Continued on next page

IRIS ALIFAX TEST 1 ESR ANALYZER

QUALITY CONTROL

Continued..

| Step | Action |
| --- | --- |
| 2. | Press KEY 6 (Electronic Calibration) to start the Latex Control Procedure. NOTE: If the instrument needs a washing, a message will be displayed “Make a washing”. Follow the instructions described above “Before Analysis, #1). Otherwise, proceed to #3. |
| 3. | Press KEY 1 to start. |
| 4. | Take a rack. - Position 1 filled with at least 3 ml of distilled water. - Positions 2,3, and 4 with the three (3) Latex Control Tubes (respecting the right order of the labels). - Positions 5 and 6 filled with at least 3 ml of distilled water. |
| 5. | When “Insert rack STD” is displayed, open the door, insert the rack in the instrument and close the door. |
| 6. | After inserting rack with test tubes, analyzer will start mixing and a “Mixing” message will be displayed, which takes about one minute an half to be completed. NOTE: If “STD N X xxxx” appears instead of Mixing, it means the scanner did not read the code on one of the test tubes. In this case, insert the code manually as described in the procedure below, “Latex Codes Typed Manually”.- |
| 7. | After mixing, the Latex Controls are analyzed and displays some messages which are shown during first washing, during latex analysis and during the last test tube of washing as follows: - A1 1023 A2 970 T100 3529 - STD procedure wait - A1 1023 A2 970 T100 3529 |
| 8. | When analysis is completed, result of the latex analysis is printed and removal of rack message is displayed |
| 9. | When “Extract rack” is displayed, open the loading door and remove the rack. Close the front door. From this position, the instrument can perform analysis cycle with blood. |

________________________________________________________________________

Continued on next page

IRIS ALIFAX TEST 1 ESR ANALYZER

QUALITY CONTROL

Continued..

LATEX CODES	At the first use of a new control kit and if the scanner cannot read one or more of

TYPED 		the three (3) codes, each missed code could be typed manually.

MANUALLY

| Step | Action |
| --- | --- |
| 1. | Press Clear twice |
| 2. | Go back to Main Menu |
| 3. | Recover the rack and take note of the three (3) codes on paper. |
| 4. | Press Key 6 and restart the Control process |

_________________________________________________________________

Before mixing , the display should show a message like:

| Message | Action |
| --- | --- |
| STD N 2 | Type the complete code of the test tube marked number 2, the space between the two parts of the code will be assigned automatically. |
| STD N 3 | Type the complete code of the test tube marked number 3, the space between the two parts of the code will be assigned automatically. |
| STD N 4 | Type the complete code of the test tube marked number 4, the space between the two parts of the code will be assigned automatically. |

___________________________________________________________

If the scanner cannot read one or more of the three (3) codes during the second or the remaining four latex processes, the display will report the code stored at the first Control process:

| Message | Action |
| --- | --- |
| STD N 2 Memo 1534 xxxxx | Press Enter and the stored code will be confirmed automatically. |
| STD N 3 Memo 1534 xxxxx | Press Enter and the stored code will be confirmed automatically. |
| STD N 4 Memo 1534 xxxxx | Press Enter and the stored code will be confirmed automatically. |

____________________________________________________________

Continued on next page

IRIS ALIFAX TEST 1 ESR ANALYZER

QUALITY CONTROL

Continued..

REVIEWING		The result print-out will include the following information:

RESULT

PRINT-OUT

- Date and time of washing – confirms the correct functioning of the photometer.

- ESR values and ranges the internal scanner reads from the labels at the beginning of the Control process.

- Instrument type (Test 1)

- Software version

- Serial Number

- Date and Time of Latex Control

- The parameter related to the analyzer gains (FDF, BoosterY, MFACT).

- The values given from the latex controls, which should tay within the expected ranges – the suitable reference values could vary from Lot to Lot.  Refer to the IFU’s to present in the kit box.

##### If controls are within the acceptable limits, proceed with patient testing.  If control result values exceeds the allowable limits, corrective action must performed and recorded before analyzing patient samples.

____________________________________________________________

Controlled	The controlled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
| LGM 2022 | Quality Manual |
| LGM 8000 | Safety Practices |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |
| LGM 8006 | Handling of Regular and Infectious Waste |

Continued on next page

IRIS ALIFAX TEST 1 ESR ANALYZER

QUALITY CONTROL

Continued..

| LGM 8007 | Cleaning Work Areas |
| --- | --- |
| LGM 8010 | Hand washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |

__________________________________________________________________

##### NON			The Non Controlled Documents are as follows:

CONTROLLED

DOCUMENTS

| Document No. | Name of Documents |
| --- | --- |
|  | Iris Alifax Test 1 Analyzer Operator’s Manual |

Author			Marianne Muyrong, CLS

IRIS ALIFAX TEST 1 ESR ANALYZER

QUALITY CONTROL

Continued..

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Julie Toti, CLS, MS Manager, Clinical Laboratory – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

Continued on next page

IRIS ALIFAX TEST 1 ESR ANALYZER

QUALITY CONTROL

Continued..

HISTORY PAGE

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Laboratory Director review/date | Laboratory Manager review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New | . |  |  |  |  |