#### I CHEM VELOCITY SAMPLE PROCESSING

| Purpose | The iChem VELOCITY is an automated urine chemistry system that performs measurements of urine physical and chemical constituents utilizing test strips read by wavelength reflectance, and specific gravity using the refractive index. ____________________________________________________________ |
| --- | --- |
| Scope Policy | The iChem VELOCITY designed to analyze urine, samples. To analyze samples, the appropriate reagents, reference materials, and QC samples are required. ____________________________________________________________ Quality control must be processed and be acceptable every 24 hours before sample results can be released. |

| Specimen | Acceptable Urine Specimen - Freshly voided clean catch in sterile cup or BD Vacutainer® Urine transport tube from MOB. - Catheterized - Bladder tap Specimen Rejection Criteria - Grossly bloody or highly colored urines should be rejected for the chemical analysis, but reflexed for the microscopic analysis. - Reject urines that exceed stability limits - Reject urines that are unlabeled or mislabeled. Minimum Volume Requirement - Two milliliters (2 cc) Specimen Stability - </=2 hours at room temperature - 0 to 24 hours, refrigerated ( 2 – 8 C) |
| --- | --- |
|  | Process Should be tested as soon as possible or within 2 (1) hours after voiding, if unable to test or delivered within recommended time, refrigerate immediately up to 24 hours and let it return to room temperature before testing. Note: Unpreserved urine at room temperature may undergo pH changes due to microbial proliferation, which may interfere with protein determination. |

- Continued on next page

#### I CHEM VELOCITY SAMPLE PROCESSING

| ##### Reagents and/or Media | Description | Vendor | Storage |
| --- | --- | --- | --- |
|  | iChem VELOCITY Urine Chemistry Strips | Beckman Coulter | Room temperature |
|  | iChem wash solution | Beckman Coulter | Room temperature |
|  | IRISSpec CA/CB/CC controls | Beckman Coulter | Refrigerated |
|  | iChem Cal/Chek Kit | Beckman Coulter | Refirgerated |

| ##### Materials and supplies | - Sample Racks - Control Rack - CalChek rack - Test tubes |
| --- | --- |

- Continued on next page

#### I CHEM VELOCITY SAMPLE PROCESSING

___________________________________________________________

| ##### Safety or Special Safety Precautions | All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety and the safety of others and adhering to all departmental and medical center safety policies and procedures. - For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). - For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. - For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing. - For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. - For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste. - For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. - For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. - For proper storage and disposal of chemical hazardous waste, see LGM 8012. . |
| --- | --- |

| ##### Definitions | IChem VELOCITY utilizes wavelength reflectance and refractive index. |
| --- | --- |

- Continued on next page

#### I CHEM VELOCITY SAMPLE PROCESSING

| ##### Procedure: | - Follow instructions below for sample handling - 1; - Use fresh urine specimen collected in clean and/or sterile container. If specimen is not tested within 2 hours specimen must be refrigerated. Specimen should always be tested at room temperature. If specimen is outside room temperature range the specific gravity may be inaccurate. Allow all refrigerated specimen to come to room temperature before testing. - Note: Samples with gross hematuria and or dense or viscous specimens must be tested using manual methods. - 2; - Transfer 3.0 ml (minimum of 2.0 ml)of well mixed urine into the appropriately barcode labeled tube .Place the barcode labeled tubes into sample racks. The following barcodes are accepted Code 128, Code 39, Codabar and Interleaved 2 of 5. Make sure to place tubes with barcode facing directly out.. - Follow the steps bleow to test urine samples. |  |
| --- | --- | --- |
|  | Step | Action |  |
|  | 1 | Log on, type the user name in the identifier field or use drop down browser on the right side of the field to display and select the user name. Type the user password in the Password field. Click OK to logon. |  |
|  | 2 | Ensure that sufficient supplies and consumable are available to complete the anticipated workload. |  |
|  | 3 | Verify that QC has been run if not run QC prior to any patient testing. Run the QC on the I Chem Velocity following the QC procedure. Note: If instrument has been off for more than 6 hours allow instrument to warm up for 50 minutes. If Qc run fails, the system will not allow the operator to run specimens. |  |
|  | 4 | After acceptable QC place sample racks containing bar-coded specimens on the right side of iChem Velocity sampler |  |
|  | 5 | If the iChem Velocity is in Standby (green light lit), press the START button. |  |
|  | 6 | If iChem Velocity is in Measure (blue light lit), block the sensor at the front of the Sampler (nearest to operator), and the rack will move to the sampling position automatically. |  |

- Continued on next page

#### I CHEM VELOCITY SAMPLE PROCESSING

___________________________________________________________

| Limitations | See below the list of conditions that may affect results and require further review |  |  |
| --- | --- | --- | --- |

| Intereferences | High levels of ascorbic acid (Vitamin C) may cause falsely lower results for glucose and blood. If clinician suspects high intake of vitamin C, suggest repeat of urinalysis after vitamin C intake has been discontinued at physician discretion |
| --- | --- |
| Hematuria | Gross hematuria may cause incorrect results in subsequent samples. Do not test specimens exhibiting gross hematuria |
| Color Intereference | Substances that cause abnormal urine color such as drugs containing azo dyes may alter accuracy of results overall |

- Continued on next page

#### I CHEM VELOCITY SAMPLE PROCESSING

___________________________________________________________

Controlled

documents            The controlled documents are as follows:

___________________________________________________________

| Document No | Name of Document |
| --- | --- |
| LGM 2022 | Quality Manual |
| LGM 8000 | Safety Practices |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |
| LGM 8006 | Handling of Regular and Infectious Waste |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand-washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |
| LUM253 | iChem Velocity QC |
| LGM2030 | Criteria for an unacceptable specimen |

Non-controlled  The non-controlled documents are as follows:

Documents              \\

|  | Document Name |
| --- | --- |
| 1 | IChem VELOCITY Operator’s Manual, 300-4449 English Rev E 10/2008 |

Author

Maaret Toprakci, MS, MT(ASCP)

#### I CHEM VELOCITY SAMPLE PROCESSING

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Maaret Toprakci, CLS, MS Area Laboratory Manager – Los Angeles Medical Center |  |
| Julie Toti, CLS, MS Laboratory Director – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

- Continued on next page

#### I CHEM VELOCITY SAMPLE PROCESSING

___________________________________________________________

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |