FilmArray® Respiratory Panel (RP) System Preventive Maintenance and Decontamination

| Daily Maintenance | Follow the steps below to perform daily maintenance. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1 | - FilmArray® Pouch Loading Stations should be rotated and cleaned at the end of each shift. - Submerge the FilmArray® Pouch Loading Station (used from the previous shift) in freshly prepared 10% bleach for 15 minutes. - Replace bleach with water and submerge the station in water to rinse. - Repeat the water rinse two additional times. Wipe dry. - Discard paper towel in a biohazardous waste container. |
| 2 | - Wipe down the surface of the FilmArray Instrument, including the sample chamber with freshly prepared 10% bleach followed by a water rinse. Alert! Thoroughly rinse with water so as to prevent the button in sample chamber from sticking. - Wipe dry the sample chamber with a paper towel. - Discard paper towel in a biohazardous waste container. |

| As Needed Maintenance | Follow the instructions below to perform “as needed” maintenance. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1 | Periodically, use a lens cloth and lens cleaner to clean the lens of the barcode reader. |

Continued on next page

FilmArray® Respiratory Panel (RP) System Preventive Maintenance and Decontamination

| Decontamination of Instrument | When there’s a spill, follow the steps below to perform decontamination of contaminated areas of the FilmArray® instrument. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1 | Put on clean PPE (gloves, lab coat and face shield). |
| 2 | Ensure that no one uses contaminated instrument. |
| 3 | - Wet a paper towel with 10% bleach (1 part bleach to 9 parts water) and wipe the affected areas. If…; Then…. sample chamber; wipe the inner sample chamber and under the lid.. instrument exterior; wipe all exterior surfaces of the instrument, including the bottom. - Change gloves |
| 4 | Repeat Step 3 twice with fresh paper towels for a total of three (3) bleach wipes. |
| 5 | Change gloves. Wet a paper towel with distilled water and wipe sample chamber. |
| 6 | Repeat the water rinse with a fresh paper towel. |
| 7 | Alert! Run a UTM sample to ensure proper function of the instrument (if the pouch leaked inside the instrument). |

| Decontamination of Other Areas | When there’s a spill, follow the steps below to perform decontamination on contaminated bench counters. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1 | Put on a clean lab coat and gloves. |
| 2 | Spray the 10% bleach solution on the area that may have been contaminated. Let it stand for 5 min. |
| 3 | Wipe the area with a clean paper towel. Change gloves. |
| 4 | Repeat Steps 2 and 3 twice, for a total of three wipes. |
| 5 | Change gloves. Spray the area with distilled water. |
| 6 | Wipe the area dry with a new paper towel. Change gloves. |
| 7 | Spray the area with DNA-Away. Change gloves. |
| 8 | Rinse the area by spraying it with distilled water and wiping it dry. |

Continued on next page