#### Beckman Coulter Unicel DxH 800 System

#### Changing Control Lot Numbers

| - Policy | When a new lot number of controls are received, the new information must be entered into the Beckman Coulter DxH800 instrument workstation. |
| --- | --- |

| - Safety | All reagents and controls should be handled as though capable of transmitting infectious diseases. Wear appropriate personal protective equipment when running patient samples or performing scheduled maintenance. Refer to Laboratory Policy and Procedure Manual Safety Section SAF100. |
| --- | --- |

| - Setting up the Controls | Follow the steps below to enter the information for the control files into the DxH 800 workstation. |
| --- | --- |

| A.1 NEW BECKMAN COULTER CONTROL FROM BARCODE |  |
| --- | --- |
| Step | Action |
| 1. | System Status screen; click QUALITY CONTROL ICON |
| 2. | Select the "More Options" from Local Navigation Icon; select QC setup. |
| 3. | Select “New Control from Bar Code.” |
| 4. | “Waiting for 2D bar code to be scanned from assay sheet” will pop up. |
| 5. | Scan bar code from the new lot number insert. |
| 6. | Click on Auto Stop |
| 7. | Select the "OK" to save. |

| A.2 NEW MANUAL ENTRY (Beckman Coulter Control): |  |
| --- | --- |
| Step | Action |
| 1. | System Status screen; click QUALITY CONTROL ICON |
| 2. | Select the "More Options" from Local Navigation Icon; select QC setup. |
| 3. | Select New Manual Entry to display the New Manual Entry Control dialog box |
| 4. | Select a type from the Type drop-down list and a control level from the Level drop-down list. |
| 5. | Select New to dislay the Create Control dialog box. NOTE: If incorrect level, source, or type, the control file must be deleted and the control must be entered again. |

#### Beckman Coulter Unicel DxH 800 System

#### Changing Control Lot Numbers

| 6. | Enter a Lot Number and select an Expiration Date from the drop-down list. CAUTION: Do not use hyphens, it will cause the control results to be stored as patient results. |
| --- | --- |
| 7. | Enter the Target and Limit values for the parameters, or choose from the Assigned Target and Expected Limit options available on the screen. Click Auto Stop for Control Settings |
| 8. | Select Save to save the control settings. |

| - Checking new lot numbers | Follow the steps below to test the new lot numbers against the old lot numbers. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Analyze the new lot of controls once and compare the values to the package insert. Results for all three levels of control should fall within the 5% of the expected range of the package insert. If it is not, consult with the supervisor. |
| 2. | The new controls are analyzed a minimum of 10 runs on different shifts/days. At the end of this time period, a new target value (mean) is established. |
| 3. | Change the assay mean to the laboratory’s established mean. |
| 4. | After running the control 20 times, determine if the mean value calculated previously is still running the same. If the control value is running above or below the target from the insert, recalculate and change the mean value. The mean value should still be within the 5% of the expected range of the package insert. Evaluate carefully and make sure the change is not an instrument problem. |
| 5. | Standard deviations in use are calculated from historical data/Lab Limits. |

#### Beckman Coulter Unicel DxH 800 System

#### Changing Control Lot Numbers

Continued,

______________________________________________________________________________

| - Handling Old Lot Data | Follow the steps below for the outgoing lot number of controls once the new lot number has been accepted for use: Step; Action. 1.; Print the data and charts for the outgoing lot.. 2.; Click the Quality Control Icon, Click the printer Icon on the upper right of the window.. 3.; Under QC Report Window, Choose Report type as Summary Report. Select outgoing lot numbers of controls. Select Full page under Graphical Reports. Click Print. 3.; File all printouts on DXH QC Summary Report Binders. |
| --- | --- |

#### _________________________________________________

- Export - IQAP Data

| Step | Action |
| --- | --- |
| 1. | Review control files. |
| 2. | From the Control Setup Screen, highlight a control level or levels and then select the Export IQAP button. |
| 3. | Under Control Data to Export, choose Selected DxH Control Lots. |
| 4. | Under Data Export Grouping, choose Use Single Shift |
| 5. | Under Destination, Choose Transfer to BCI, Click Start to upload data. |
| 6. | eIQAP reports will available online 1-2 Days from time of upload |

When ready to submit control data to IQAP:

#### Beckman Coulter Unicel DxH 800 System

#### Changing Control Lot Numbers

Continued,

______________________________________________________________________________

| eIQAP Step | Action |
| --- | --- |
| 1. | Go to Beckmancoulter.com |
| 2. | Select Hematology eIQAP |
| 3. | Login: 4867Sunset Password: Kaiser123 |
| 4. | Choose Instrument: DxH 800 [06886-1-B1] – DxH 800 #4 DxH 800 [06886-1-B2] – DxH 800 #2 DxH 800 [06886-1-B3] – DxH 800 #3 |
| 5. | Choose Reports to Print. |
| 6. | File all Reports to DXH eIQAP Binders |

Download of Reports

Reviewed and approved by:

Continued on next page

| SIGNATURE | DATE |
| --- | --- |
| Julie Toti, CLS, MS Manager, Clinical Laboratory – Los Angeles Medical Center |  |
| Joseph C. Thompson, M.D. Director, Clinical Laboratory – Los Angeles Medical Center |  |

#### Beckman Coulter Unicel DxH 800 System

#### Changing Control Lot Numbers

Continued,

______________________________________________________________________________

#### Beckman Coulter Unicel DxH 800 System

#### Changing Control Lot Numbers

Continued,

______________________________________________________________________________

#### HISTORY PAGE

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Laboratory Director review/date | Laboratory Manager review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  |  |  |  |  |