Beckman Coulter Unicel DxH 800 System

Quality Control

##### PURPOSE	The purpose is to provide instructions on running the Quality Control materials before patient sample analysis.

________________________________________________________________________

##### POLICY	The 6C Cell Control and Retic Control must be analyzed at the beginning of each eight (8) hours of operation plus or minus 15 minutes and be within the established limits before patient testing can be started.  The scheduled Quality Control testing as follows:

Day Shift - 06:30 AM  (+/- 15 mins)

Evening Shift - 2:30 PM (+/- 15 mins)

Night Shift – 10:30 PM  (+/- 15 mins)

##### _________________________________________________________________

##### SCOPE	The 6C Cell Control and Retic Control must be run after any scheduled or unscheduled preventive maintenance or repair, after calibration, or calibration verification.  The controls are run as a patient, using the same reagents, and performed by the same personnel testing patient samples. Levey Jennings Graph must be reviewed for each parameter at every QC run. When quality control tolerance limits are exceeded, corrective action must performed and recorded before analyzing patient samples.

The Latron CP-X Control must be analyzed after start-up on a daily basis.  Latron CP-X Control also monitors the stability of the electrical processing and the fluidic flow rate systems used to measure the volume, conductivity and light scatter.

Workplace	All laboratory employees are expected to maintain a safe working environment

Safety		and an injury-free workplace.  Laboratory employees are responsible for their

own safety, and the safety of others and adhering to all departmental and medical

center safety policies and procedures.

- For standard precautions and safety practices in the laboratory;  see LGM

8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids.

- For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing.

- For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves.

- For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and bio hazardous waste.

- For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued,

- For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- For proper storage and disposal of chemical hazardous waste, see LGM 8012.

| Parameters run on Controls | 6C Cell Control - WBC, RBC, HGB, HCT, MCV, MCH, MCHC,RDW, RDW-SD, PLT, MPV, NE%, LY%, MO%, EO%, BA%, NRBC, NE#, LY#, MO#, EO#, BA#, NRBC# Retic- X Control - RET |
| --- | --- |

MATERIALS	COULTER® 6C Cell Control

COULTER® Retic-X Cell Control

COULTER® Latron CP-X Control

Specimen Casettes

Gloves

Specimen Cassettes

| QC Setup | Follow procedure for QC Setup on LHM 299.16 Beckman Coulter DxH800 Changing Control Lot Numbers |
| --- | --- |

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued,

- Running a Quality Control

A.1.	COULTER® Latron CP-X Control

| Step | Action |
| --- | --- |
| 1 | Note: Bring Coulter Latron CP-X Control to Ambient Temperature before use. The Control is used when performing Daily Checks. Gently mix at room temperature by inversion of five (5) to eight (8) times. AVOID FOAMING. |
| 2 | Place the controls into a sample cassette ensuring the barcode is viewable. |
| 3 | Load the cassette into the input buffer. |
| 4 | After the running the control materials, and results are within acceptable limits, store in room temperature. |

A.2.	COULTER® 6C Cell Control

| Step | Action |
| --- | --- |
| 1 | Remove control tube(s) from the refrigerator and warm at Ambient Temperature for 10-15 minutes. |
| 2. | After warming, mix by hand as follows: - Roll the tube slowly between the palms of the hands eight (8) times in an upright position. - Invert the tube and slowly roll it between the palms eight (8) times. - Gently invert the tube eight (8) times Note: Perform 2a to 2c two (2) times. Do Not Over Mix. Do Not use any type of mechanical mixer. |
| 3 | Put instrument Online by selecting the Start Button from the upper right, on the system icons. Place the controls into a sample cassette ensuring the barcode is viewable. |
| 4 | Place the controls into a sample cassette ensuring the barcode is viewable. |
| 5 | Load the cassette into the input buffer. |
| 6 | After the running the control materials, and results are within acceptable limits, return the control tube(s) to the refrigerator within 30 minutes. |

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued..

A.3.	COULTER® Retic-X Cell Controls

| Step | Action |
| --- | --- |
| 1 | Remove control tube(s) from the refrigerator and warm at Ambient Temperature for 10-15 minutes. |
| 2. | After warming, mix by hand as follows: - Roll the tube slowly between the palms of the hands eight (8) times in an upright position. - Invert the tube and slowly roll it between the palms eight (8) times. - Gently invert the tube eight (8) times Note: Perform 2a to 2c ONCE ONLY. Do Not Over Mix. Do Not use any type of mechanical mixer. |
| 3 | Put instrument Online by selecting the Start Button from the upper right, on the system icons. Place the controls into a sample cassette ensuring the barcode is viewable. |
| 4 | Place the controls into a sample cassette ensuring the barcode is viewable. |
| 5 | Load the cassette into the input buffer. |
| 6 | After the running the control materials, and results are within acceptable limits, return the control tube(s) to the refrigerator within 30 minutes. |

Continued on next page

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued..

- Reviewing Quality Control Results

B.1. COULTER® 6C Cell Control and Retic-X Control Results

| Step | Action |
| --- | --- |
| 1. | Log in to your DXH instrument profile. |
| 2. | Select the Quality Control Status Icon. NOTE : If any of the control runs had out-of-range values, a dialogue box opens with a list of the those controls. If everything was within range, the Quality Control (Data View) screen opens with the last control that was viewed. |
| 3. | To review a different control/level, use the Select Control button from the local navigation bar. |
| 4. | Choose the control/level for review, hightlight and select OK. |
| 5. | The most recent run will be on the right side, column is highlighted blue. LJ graph must be reviewed for each parameter per QC level at each QC run to check for any Shifts and Trends. |
| 6. | If all results are within acceptable limits, click “Reviewed by” to document that QC had been reviewed. |
| 7. | LJ graphs can be viewed by selecting graph button from the local navigation bar. (To return to the data screen, select view data button). |
| 8. | For results outside the acceptable limits, result values will be in red color. Select the “Reviewed By” to document that unacceptable results were reviewed. |
| 8. | Rules applied are: 13S, 2of32S |
| 10. | Select the Comment button from the local navigation bar to document any corrective action performed. Enter Rule(s) violated if 13S, 2of32S, and if there is Shifts, or Trends occurring on a parameter of that level. If 2 levels are having issues, make sure to put a comment on both levels for corrective action performed. You may add or modify comments after saving them. (Once saved, a “yellow page” icon will appear to denote a comment is available to view). |

Continued on next page

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued..

| 11. | Follow QC Criteria table for corrective action NOTE: “Exclude button” is available at the bottom of the column, only exclude bad runs such as R-Flagged runs and Partial aspirate/vote out runs. |
| --- | --- |

| QC Criteria | Definition | ACTION |
| --- | --- | --- |
| 13s | Indicates that control limits are set as the mean plus 3s and the mean minus 3s. A run is rejected when a single control measurement exceeds the mean plus 3SD or the mean minus 3SD control limit. | Repeat run on QC level greater than 3SD. Do further investigation if run is still out. |
| 2of32s | The run is rejected when 2 of 3 Controls exceed the mean value +2SD or the mean –2SD limits. It should be applied within and across runs. This rule is violated within the run when two consecutive control values of three controls exceed the “same” ( mean 2SD) limit. The rule is violated across runs when the previous value for a particular control level exceeds the "same" (mean 2S) limit. | Repeat run on QC levels that violated this rule. Do further investigation if run is still out. |

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued..

A trend in a quality control chart refers to a series of 5 to 7 daily control values that contribute either to increase or decrease from the mean value.

A shift in a quality control chart is when 5 to 7 or more daily control values for a period of days distribute themselves above or below the mean value, but are maintaining a constant level.

Trends and Shifts are not used as criteria to reject a quality control run as long the quality control run conforms to the primary QC criteria for the particular testing aforementioned above. Please notify the Laboratory Manager to initiate a corrective action if a problem exists in the procedure or instrumentation. In both cases, if a problem exists, immediate action must be taken to resolve the problem and documentation must indicate the reason for the shift or trend and the corrective action taken. Comments for corrective action are entered on the last QC Run of the level having issues in the instrument.

B.2. COULTER® Latron CP-X results

| Step | Action |
| --- | --- |
| 1. | Select the Quality Control Status Icon. |
| 3. | Choose Coulter Latron CP-X control for review, hightlight and select OK. |
| 4. | The most recent run will be on the right side, column is highlighted blue. Values outside the expected range are in red. |
| 5. | Examine the Mean/Mode Channel and Coeficient of Variation results for the volume, conductivity and various light scatter measurements. Compare the results to the Table of Expected Results. |
| 6. | If run is unacceptable, make sure control material does not have any bubbles, then proceed with re-running the Latron CP-X Control. |
| 7. | Select the Comment button from the local navigation bar to document any corrective action performed. |
| 8. | After re-running Latron CP-X Control and results are still out, inform Manager, and then call Technical Support for assistance. |

_________________________________________________________________

Continued on next page

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued..

Controlled	The controlled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
| LGM 2022 | Quality Control Policy |
| LGM 8000 | Safety Practices |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |
| LGM 8006 | Handling of Regular and Infectious Waste |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |

_________________________________________________________________

##### NON			The Non Controlled Documents are as follows:

CONTROLLED

DOCUMENTS

| Document No. | Name of Documents |
| --- | --- |
|  | Beckman Unicel DxH Coulter Operator’s Manual |

_________________________________________________________________

Author			Alvin Castillo, CLS

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued..

__________________________________________________________________

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Alvin Castillo, CLS Supervisor, Clinical Laboratory- Los Angeles Medical Center |  |
| Jocelyn Javier, CLS Quality Manager, Clinical Laboratory – Los Angeles Medical Center |  |
| Julie Toti, CLS, MS Laboratory Director, Clinical Laboratory – Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Medical Director, Clinical Laboratory – Los Angeles Medical Center |  |

Beckman Coulter Unicel DxH 800 System

Quality Control

Continued..

## HISTORY PAGE

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Laboratory Director review/date | Laboratory Manager review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New | . |  |  |  |  |