| Purpose | - This document will define the procedure for the use of performing a manual complete semen analysis which will indicate the presence of sperm from a collected fresh sample or a post-vasectomy sample. |
| --- | --- |

| ##### Policy | - Patients are instructed to bring specimens to the laboratory within 1 hour of collection. The specimen must be maintained at room temperature and not exposed to heat or cold. - Semen analysis is performed at room temperature between 680-75.20 F (200 – 240C). Temperature deviations are noted in report. - Semen Analysis: Manual method is performed when Semen Analysis automated instrument is non-operational, questionable instrument results that need reviewing, or low volume samples < 300 µL. |
| --- | --- |

| Workplace Safety | All laboratory employees are expected to maintain a safe working environment Safety and an injury-free workplace. Laboratory employees are responsible for their own safety, and the safety of others and adhering to all departmental and medical center safety policies and procedures. • For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). • For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. • For proper hand-washing, see LGM 8010, specifically, not limited to, proper hand-washing. • For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. • For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste. • For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas. • For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. • For proper storage and disposal of chemical hazardous waste, see LGM 8012.All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures. |
| --- | --- |
| Specimen | - Fresh – samples collected within an hour of testing and not enriched, diluted or treated. - Post-vasectomy – fresh samples collected within an hour of testing and designated as post-vasectomy specimens. - The patient’s physician or the laboratory provides the collection instructions and the sterile plastic container. - Specimen must be completely labeled with patient’s name and medical record number. - Specimen must be protected from extremes of temperature (less than 680F or more than 1040F) during transport to the laboratory. - Specimen must be delivered to the laboratory within 1 hour of collection and must reach the testing laboratory within 30 minutes for an accurate evaluation of sperm motility. |

| ##### Equipment and Reagents | Assemble the following supplies before proceeding: - Incyto Disposable Hemocytometer - Semen Diluting Fluid (Qwik Earle’s Solution) - Qwik Chek Liquefaction Kit - microscope slides - glass coverslip 22 X 22 mm - IRIS CA & CB Urinalysis QC & Qwik Check Test Strips |
| --- | --- |

| ##### Quality control | - QC Qwik Check Test Strips using Biorad Urine Controls: IRIS CA = <1 M/ml Leukocytes (AM SHIFT) IRIS CB = ≥1 M/ml Leukocytes (PM SHIFT) - QC is run once every 24 hours when you receive a sample |
| --- | --- |

Macroscopic Examination:

| Step | Action |
| --- | --- |
| 1 | Allow the specimen to liquefy. Swirl semen in container to determine if coagulum has liquefied. A liquefied specimen will take the shape of the container. Note: - Normal semen sample liquefies within 60 minutes at room temperature. Continuous gentle mixing or rotation of specimen during liquefaction may reduce errors in determining sperm concentration. Failure of specimen to liquefy within 60 minutes must be recorded in the report. When samples do not liquefy within 60 minutes, extend incubation time and pipette the specimen repeatedly with a sterile pipette. If this method fails, use of an enzymatic treatment- Qwik Check Liquefaction Kit may be necessary. Use of these manipulations must be recorded in the report. |
| 2 | Mix liquefied specimen well. |
| 3 | Pour specimen into a graduated plastic centrifuge tube and determine the volume to the nearest 0.1 mL. Record volume in mL. |
| 4 | Determine the viscosity of the liquefied sample. - Draw the specimen into a disposable pipette and allow to drop back into container by gravity. Normal; Specimen slightly mucoid, but will fall in separate drops.. Decreased; Specimen watery, not viscous or mucoid.. Increased; Specimen very mucoid, will not fall in separate drops. The drop will form a thread more than 2 cm long.. Note: - High viscosity can interfere with determination of sperm motility and concentration. The methods to reduce viscosity are the same as those for delayed liquefaction. |
| 5 | Quantitate WBC’s using Qwik Check test strips or comparable. |
| 6 | Report as WBC </= 1M/ml or WBC >1 M/ml |

Microscopic Examination

| Step | Action |
| --- | --- |
| 1 | Using an pipette deliver 10 L of well mixed liquefied semen sample onto a clean glass slide and cover sample with a 22 mm X 22 mm glass cover slip. The freshly made wet preparation is left to stabilize for approximately one minute. Note: - Examination is performed at room temperature between 68–75.2 0F. Record examination temperature in sperm assay log. |
| 2 | Perform initial evaluation at 10X. Scan slide for mucus strand formation, sperm aggregation, and evenness of spread of sperm on slide. If the number of sperm per visual field varies considerably, sample is not homogeneous. Remix the sample thoroughly and repeat exam. Note: - Lack of homogeneity may also be due to abnormal consistency, abnormal liquefaction, aggregation of sperm in mucus threads, or from sperm agglutination. These abnormalities must be mentioned in the report. - If no spermatozoa are seen during the initial examination at 10X and at 40X, centrifuge entire sample 2000 RPM for 3 minutes. A complete and systematic search of all the re-suspended precipitate must be done. |
| 3 | At 40X scan slide for motility and sperm concentration. |

Assessment of Sperm Motility

|  | Assess sperm motility: At least five microscopic fields are assessed in a systematic way to classify 200 spermatozoa. Examination is performed within one hour of specimen collection, once complete liquefaction has occurred. 0; Immotility in percentage. 1+ 2+; Non- progressive motility in percentage. 3+; Slow progressive motility in percentage. 4+; Rapid progressive motility in percentage. |
| --- | --- |

Assessment of Sperm Concentration:

| Step | Action |
| --- | --- |
| 1 | Mix semen sample gently and thoroughly. |
| 2 | Dilute semen with semen diluting fluid either 1:10 or 1:100 . Use serial dilution. For 1/10 dilution label a tube 1:10 with patient name and accession number pipette 0.1 ml of well mixed semen into the tube and add 0.9 ml semen diluting fluid, mix well. To make 1:100 dilution take 0.1 ml of 1:10 dilution pipette it into a tube labeled 1:100 patient name and accession number add into this 0.9 ml semen diluting fluid mix well. Note: For low volume specimens dilutions can be made using smaller volume of semen. |
| 3 | Load 10 µl of diluted sample on both sides of the hemocytometer either a 1:10 or 1:100 dilution. Allow the hemacytometer to settle for 1-2 minutes. |
| 4 | Using the high dry objective, count the sperm in the 5 small squares(4 small corner squares and one small middle square) in the center square of the chamber (RBC counting area) in each chamber and average the count. Note: Both sides of the chamber are counted and counts must be within 10%. Rejected specimens must be mixed and recounted. If still not acceptable, new dilutions should be set-up and recounted. |
| 5 | Calculate the sperm count according to the type of counting chamber using the following formula: In cyto c- chip Hemocytometer: Cells in 5 small squares x 5 x dilution factor x 104 (volume factor) = sperm/mL Example: Count 1= 55 sperm seen Count 2=50 sperm seen Average count = 53 sperm seen (9% difference) 53 x 5 x 10 x 10000 = 26.5M/mL |
| 6 | If a count is low perform the sperm count with no dilution. |
| 7 | For Sperm Concentration of <2M/ml(lowest reportable range) – Enter under result comment if No sperm/hpf or Rare sperm/hpf. |

Assessment of Sperm Morphology:

| Step | Action |
| --- | --- |
| 1 | Evaluate sample for Normal forms present. Strict criteria are applied in assessing the morphological normality of the spermatozoon. For spermatozoon to be considered normal, the sperm head, neck, midpiece, and tail must be normal. Head must be oval in shape. - Mix the semen well. - Place a drop of semen on a slide and with a second slide drag the sample along the surface. Make 2 smears. - Air dry smear, label, and stain on the Hematek slide stainer. - Using oil immersion objective, evaluate 200 spermatozoa in slide area where sperm are evenly distributed. Spermatozoa are recorded as either normal or abnormal. - Report % normal forms. Head Defects; Neck & Midpiece; Tail Defects. - Large - Small - Tapered - Pyriform - Round - Amorphous head - Vacuolated heads (>20% head area occupied by unstained vacuolar areas) - Heads with small acrosomal ara (<40% of head area) - Double heads - Or any combination of the above.; - Bent neck (the neck and tail form an angle of >90% to the long axis of the head) - Asymmetrical insertion of the midpiece into the head - Thick or irregular midpiece - Abnormally thin midpiece - Or any combination of the above.; - Short - Multiple - Hairpin - Broken tails - Bent tails (>900) - Irregular width - Coiled tails - Or any combination of the above.. |

Resulting in Cerner:

| Step | Action |
| --- | --- |
| 1 | Access Cerner ARE. |
| 2 | Parameter; Reportable Range. Analysis Time; Enter Military Time. Volume; Record Volume in milliliters to the nearest 0.1 mL. % Normal Morphology; Enter % of normal forms according to WHO 3rd Edition. WBC Concentration; Enter 0 for <=1 M/mL (Normal) or enter 1 for >1 M/mL (Abnormal) leukocytes using the QwikCheck Test Strips. Spm Dens. Ins; Sperm concentration in million/mL. Record Motility Immotile; 0; No movement. Non Prog. Motility; 1+; sperm moving, but with no to very slight forward motion. 2+; sperm moving aimlessly, but with more direct slow forward motion. Slow Prog. Motility; 3+; sperm moving at moderate speed in a straight forward motion. Rapid Prog. Motility; 4+; sperm moving at high speed in a straight forward motion. Total Motility count; Cerner will automatically calculate and auto populate. Equation calculated by Cerner: (sperm volume*sperm conc.)*(motility 3 pct. + motility 4 pct)/100.. |
| 3 | Verify Results in Cerner. |

| Reference Ranges | Parameter; Reference Range. Volume; >1.9 mL. Normal Morphology; >29 %. WBC Concentration; <= 1 M/mL. Sperm Concentration; >15 M/mL. Total Progressive Count [Motility (3+) + (4+)]; >39 %. Immotile; None. Non Progressive Motility [ (1+) + (2+)]; None. Slow Progressive Motility (3+); None. Rapid Progressive Motility (4+); None. |
| --- | --- |

| Procedural Notes | - When the semen viscosity is greatly increased, counting will be facilitated by adding mucolytic agent such as Qwik Check Liquefaction. If a significant amount is added, a dilution correction factor must be used to calculate the count. - Another option for liquefying very viscous semen is by drawing up the sample up and down through a needle with a syringe to break up the coagulum. - Sperm are susceptible to sudden temperature changes that can affect motility. The specimen must be maintained at room temperature (680F-75.20F). - All glass and plastic ware must be clean to avoid contaminants that can affect sperm motility. - In normal semen, the majority of cells are mature sperm. Other typical cells are epithelial cells, immature germ cells, and WBCs. High percentages of these cells, however, should be reported, as should gross bacterial, RBCs, trichomonas or yeast, as a comment in Cerner at the Morphology DTA. - If patient indicated that there was a temperature problem during transportation of specimen to laboratory, the CLS must enter the information in Cerner as a comment at the exam time DTA . - Manual Methodology is performed on the following: - Motility of 0% - Viscous samples - Too many particles to give an accurate count |
| --- | --- |

| Limitations | - Some specimens may be too viscous, even after addition of a mucolytic agent, or contain too many particles to give an accurate count. In this case, attach a comment, “Unable to perform accurate sperm count due to incomplete liquefaction of the seminal fluid.” On the Sperm Count DTA. |
| --- | --- |

| Author | Alvin Castillo, CLS |
| --- | --- |

| ##### References | - WHO Laboratory Manual for the examination of human semen and sperm-cervical mucus interaction, 4TH ed.. Cambridge, UK: Cambridge University Press, 2000. - Baker, DJ.. Performing A Quality Semen Analysis in the Clinical Laboratory, MLO. December 2000. - SQA-V Gold User Guide. Version 2.49. Catalog #5109. November 2006. - V-Sperm Gold User Guide. Version 3.49. Catalog #5110. November 2006. - INCYTO C-CHIP INSERT - Qwik Check test Strips Insert - Qwik Check Liquefaction Kit Insert - Qwik Check Dilution Insert |
| --- | --- |

| ##### Controlled documents | See below for controlled documents |
| --- | --- |

| LGM8000 | Standard precautions and safety practices in the laboratory |
| --- | --- |
| LGM8005 | Universal Body Substance precautions |
| LGM8010 | Proper hand-washing |
| LGM8006 | Infection control |
| LGM8007 | Cleaning Work Areas |
| LGM8012 | Proper storage and disposal of chemical hazardous waste |
|  | Los Angeles Medical Program: Instructions for Collection of Sample for Semen Analysis |
| LHM 229 | Semen Analysis: Manual Method |
| LHM 248 | Semen Analysis: Spermalite SQAV |