- STA-R EVOLUTION CALIBRATION/VERIFICATION

- Purpose	To define calibration/verification procedures for Anti-Xa, LMWH (Low

- Molecular Weight Heparin) and D-Dimer assays in the STA-R Evolution®.

- There must be a valid Standard Curve in memory before controls and patient

- samples are analyzed.

Policy

- Calibration/Verification is performed on these assays based on the following criteria:

- Change of reagent lots for chemically or physically active or critical components

- If QC materials reflect an unusual trend or shift or are outside of the laboratory’s acceptable limits, and other means of assessing and correcting unacceptable control values fail to identify and correct the problem

- After major preventive maintenance or change of critical instrument component

- At least every 6 months

Workplace	All laboratory employees are expected to maintain a safe working environment

Safety		and an injury-free workplace.  Laboratory employees are responsible for their

own safety, and the safety of others and adhering to all departmental and medical

center safety policies and procedures.

- For standard precautions and safety practices in the laboratory;  see LGM

8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids.

- For proper hand washing, see LGM 8010, specifically, not limited to, proper hand washing.

- For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves.

- For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to, proper disposal of regular and biohazardous waste.

- For proper cleaning of work area, see LGM 8007 – Cleaning Work Areas.

- For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- STA-R EVOLUTION CALIBRATION/VERIFICATION

- For proper storage and disposal of chemical hazardous waste, see LGM 8012.

_______________________________________________________________

Procedure	 Anti- Xa and LMWH Calibration

- Materials	STA® - Owren-Koller Buffer

- STA® - Liquid Anti-Xa Assay kits (UFH and LMWH)

- STA® - Quality HNF/UFH 2 and 7 controls

- STA® - Quality HBPM/LMWH 8 and 14 controls

- STA® Multi Hep Calibrator

| Step | Action |
| --- | --- |
| 1 | For Anti-Xa calibration, reconstitute the following vial of the STA® Multi Hep Calibrator with 1.0 ml reagent grade water. - Reagent 1: STA® - Multi Hep Calibrator 0: human plasma free of heparin (see enclosed Assay Value Insert) - Reagent 2: STA® - Multi Hep Calibrator 4: human plasma containing a well defined quantity of UFH (see enclosed Assay Value Insert) - Reagent 3: STA® - Multi Hep Calibrator 7: human plasma containing a well defined quantity of LMWH (see enclosed Assay Value Insert) - Reagent 4: STA® - Multi Hep Calibrator 10: human plasma containing a well defined quantity of UFH that is greater than that of Reagent 2 (see enclosed Assay Value Insert) Let stand 30 minutes at room temperature. Swirl gently to ensure each vial is completely mixed. Reconstituted stability is 4 hours on Analyzer. |
| 2 | Request the product drawer to open by clicking the icon, select load products, bar code each vial of calibrator and place each vial of calibrator into the product drawer, R0. The assay value for each Heparin Level is transferred to the STA-R Evolution® when each calibrator is loaded and prompted for the bar code information. |
| 3 | All standard dilutions are automatically prepared by the STA-R Evolution® by diluting the STA® - Multi Hep Calibrators with STA® - Owren-Koller Buffer according to the parameters entered in the Test Set-up. |

STA-R EVOLUTION CALIBRATION/VERIFICATION

| 4 | Click on Calibration menu icon |
| --- | --- |
| 5 | Click HYBRID TEST (on the test abbreviation for which a calibration has to be run.) - Calibration screen includes current calibration curves and product information - Position for two different lots of reagent are available |
| 6 | Click on calibration icon or calibration icon (if both positions are in use, delete one calibration |
| 7 | Double click HYBRID TEST |
| 8 | Confirm by clicking on Calibrate….. |
| 9 | Select the lot # to be calibrated... (usually the one highlighted in green) |
| 10 | Confirm by clicking on * The triangle corresponding to the test to be calibrated is displayed in red. Once the test has been calibrated, the triangle is displayed in green. |
| 11 | Evaluate calibration curve for acceptability. Re-run unacceptable points if any on the calibration curve: - Click the Calibration menu icon - Single click HYBRID TEST - Calibration screen includes current calibration curves and product information - Select the calibration icon or the calibration icon - If a data point on the calibration curves needs to be repeated, double click the calibration point to rerun... - Confirm by clicking on |

STA-R EVOLUTION CALIBRATION/VERIFICATION

| 12 | Verify calibration curve is valid: Check that the correlation coefficient (r-value) of the calibration curve is 0.980 or greater. Compare the actual vs the interpolated values of the curve. If a data point difference is greater than 10% of actual vs. interpolate value, the point can be deleted or repeated. When this is done, the STA-R Evolution® will automatically re-draw the curve and re-calculate the r-value. |
| --- | --- |
| 13 | Run STA® - Quality HNF/UFH 2 and 7 controls for Anti-Xa and STA® - Quality HBPM/LMWH 8 and 14 controls for LMWH. If any control(s) is outside the expected range, the control must be repeated. If the value is still outside the expected range, a new calibration curve must be run along with fresh controls. |

D-Dimer Calibration Verification

D-Dimer Calibration Verification is performed using the CAP Calibration

Verification/Linearity kit.  The process of receiving, processing, testing,

reporting and evaluation of calibration verification is the same as

proficiency testing.  Refer to QM 5.6.4.100.

Controlled	The controlled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
| LCOM 221 | STA-R Evolution Sample Processing |
| LCOM 219 | STA-R Evolution Quality Control |
| LGM 2041 | Calibration Verification |
| QM 5.6.4.100 | Proficiency Testing |
| LGM 8000 | Safety Practices |
| LGM 8004 | Infection Control |
| LGM 8005 | Universal Body Substance Precautions |

STA-R EVOLUTION CALIBRATION/VERIFICATION

| LGM 8006 | Handling of Regular and Infectious Waste |
| --- | --- |
| LGM 8007 | Cleaning Work Areas |
| LGM 8010 | Hand washing Policy |
| LGM 8012 | Storage and disposal of Chemical Hazardous Waste |

Uncontrolled		The uncontrolled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
|  | STA-R Evolution Reference Manual |