| ##### Purpose | Iris Diagnostics recommend that the instruments remain on at all times. However, in the event of maintenance, the need to reboot, or emergency, the instrument must be shut down and powered up in a specific sequence in order to achieve correct communication between the instrument, the PC workstation of the instrument, and the laboratory information systems. |
| --- | --- |

| ##### Workplace Safety | All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures. - For standard precautions and safety practices in the laboratory; see LGM 8000, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE). - For Universal Body Substance precautions, see LGM 8005, specifically, but not limited to, exposure to body fluids. - For proper hand-washing, see LGM 8010, specifically, but not limited to, proper hand-washing. - For proper infection control, see LGM 8004, specifically, but not limited to, proper use of gloves. - For proper handling of regular and infectious waste, see LGM 8006, specifically, but not limited to proper disposal of regular and bio-hazardous waste. - For proper cleaning of work area, see LGM 8007 - Cleaning Work Areas. - For proper handling of chemicals and reagents, see the Chemical Hygiene Plan. |  |
| --- | --- | --- |

| ##### Procedure |  |
| --- | --- |

Follow the steps to Shutdown the iQ200 and Velocity

| Step | Action |
| --- | --- |
| 1 | Click the Instrument button on the top right side of the main screen of the PC workstation. |
| 2 | Click the Maintenance button located at the bottom of the Instrument screen. |
| 3 | Click the Shutdown button. The system will prompt: “Do you want to shut down the instrument?” |
| 4 | Click on the Yes button. Windows will close |
| 5 | Turn off the iChem Velocity by pressing the green button located on the left of the chemistry system. |
| 6 | Turn off the iChem Velocity power by pressing the main power switch located on the back of the chemistry system. |
| 7 | Turn off the iQ200 by pressing the Power Standby/On button at the bottom right of the instrument. |
| 8 | Turn off the iQ200 power by pressing the main main power switch located on the back of the microscopy system. |

Follow the steps to Power Up the iQ200 and Velocity

| Step | Action |
| --- | --- |
| 1 | Start the PC Module. |
| 2 | Start the iChem Velocity by pressing the “ON” switch at the back of the chemistry module. |
| 3 | Start the iQ200 by pressing the “ON” switch at the back of the microscopy module. |
| 4 | When the Instrument Screen appears on the PC workstation monitor, turn on the iQ200 by pressing the power button on the front of the microscopy module. The button light will turn green. |
| 5 | Turn on the power of the iChem velocity by pressing the Standby power switch located on the front of the instrument. The button will turn green. |
| 6 | Run Quality Control in accordance with policy. |

___________________________________________________________

Controlled

documents            The controlled documents are as follows:

| Document No | Name of Document |
| --- | --- |
| LGM 2022 | Quality Control Program |
| LUM 243 | Quality Control on IQ 200 |
| LUM 253 | iChem Velocity QC |

Non-controlled  The non-controlled documents are as follows:

Documents

|  | Document Name |
| --- | --- |
| 1 | IChem Velocity Operator’s Manual, 301-7146 English Rev B 11/2/2011 |
| 2 | iQ200 Operator’s Manual, 300-4320 English Rev C 7/2012 |

Author

Anthony Anico, CLS, MT(ASCP)

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Lab Manager. Reviewed/ Date | Med . Dir. reviewed/ date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  | AC 9/21/15 |  |  |  |

Reviewed and Approved by:

| SIGNATURE | DATE |
| --- | --- |
| Alvin Castillo, CLS Supervisor, Clinical Laboratory – Los Angeles Medical Center |  |
| Julie Toti, CLS Laboratory Director, Clinical Laboratory - Los Angeles Medical Center |  |
| Joseph Thompson, M.D. Medical Director - Los Angeles Medical Center |  |