|  |  |
| --- | --- |
| Purpose  | To provide instructions for performing qualitative Color mono agglutination test, by Sure-Vue.  |
| Policy | Quality Control for Sure Vue Color Mono Test is required at least once in 24 hours of testing.  |

|  |  |  |  |
| --- | --- | --- | --- |
|

|  |  |
| --- | --- |
| Workplace Safety | ***Refer to the safety manual for general safety requirements.*** |

 | All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures.* For standard precautions and safety practices in the laboratory; see **Safety Practices**, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).
* For Universal Body Substance precautions, see **Universal Body Substance Precautions**, specifically, but not limited to, exposure to body fluids.
* For proper hand-washing, see **Hand washing Policy**, specifically, not limited to, proper hand-washing.
* For proper infection control, see **Infection Control**, specifically, but not limited to, proper use of gloves.
* For proper handling of regular and infectious waste, see **Handling of Regular and Infectious Waste**, specifically, but not limited to, proper disposal of regular and biohazardous waste.
* For proper cleaning of work area, see **Cleaning Work Areas**.
* For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.
* For proper storage and disposal of chemical hazardous waste, see **Storage & Disposal of Chemical Hazardous Waste**. All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures.
 |

|  |  |
| --- | --- |
| Principle | The Sure-Vue Color Mono reagent is a suspension of specially treated horse red blood cells. Added coloration of the suspension facilitates the recognition of positive and negative reactions. The serum or plasma being tested is mixed on a test slide with the reagent. The appearance of dark agglutinates against a blue-green background indicates the presence of IM. If no heterophile antibodies are present, the horse cells remain unagglutinated against a green-brown background.  |

|  |  |
| --- | --- |
| Reagents | The reagents will remain stable through the expiration date, shown on the label, if stored between 2 and 8 degrees Celsius. Do not freeze.* Red Cell Reagent: Color-enhanced treated horse red blood cells suspension.
* Positive Control: Diluted positive human serum.
* Negative Control: Non-reactive diluted human serum.
 |

|  |  |
| --- | --- |
| Procedure | Follow the steps outlined below for Sure-Vue Color Mono testing |
| Sure-Vue Color Mono Testing |
| **Step** | **Action** |
| 1 | Bring all reagents and samples to room temperature (20-30 C) before testing. |
| 2 | Determine the number of separate rings that will be needed. A separate ring will be needed for each sample and control. |
| 3 | Shake the vial of Dyed, Color-Enhanced Horse Erythrocytes gently to uniformly mix the suspension. Add one free falling drop of reagent to each ring in use on the slide. |
| 4 | Using a clean pipette for each sample, place one free falling drop of undiluted specimen into center of each labeled ring.  |
| 5 | Infectious mononucleosis (IM) positive and negative controls are run with series of test specimens. Place a free falling drop of negative control serum into the center of the ring labeled negative control by squeezing the dropper. Then place a free falling drop of positive control serum into the center of the ring labeled positive control by squeezing the dropper. |
| 6 | Mix the sample or control and the reagent in each circle using a stirrer for each sample or control covering the whole surface of the circle.. |
| 7 | Gently rotate slide for ONE minute. Then allow it to remain UNDISTURBED on a flat surface for an additional one minute.  |

**Procedure,**

**Continued**

|  |  |
| --- | --- |
| **8** | Examine immediately for agglutination and record the results. Note: Strong reactions may appear within 10 seconds. |
| **9**  | Log all results in Mono Test QC and Patient Log form |
| **10** | Enter result manually in Cerner using Accession Result Entry.***Note: Follow Manual result entry protocol (click perform, review results then click verify)***. |

|  |
| --- |
| **REFERENCE FOR INTERPRETAION OF RESULTS** |
| **POSITIVE*** **Large clumping with clear blue-green background.**
* **Mod clumping with blue-green background.**
* **Small clumping with green-brown background**
 |
| **NEGATIVE*** **Roughness or not visible clumping with green-brown background**
 |

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Reference |

|  |  |
| --- | --- |
| **Document Number** | **Document Name** |
| LAMC-PPP-0123 | Safety Practices |
| LAMC-PPP-0127 | Infection Control |
| LAMC-PPP-0128 | Universal Body Substance Precaution |
| LAMC-PPP-0129 | Handling of Regular and Infectious Waste |
| LAMC-PPP-0130 | Cleaning Work Areas |
| LAMC-PPP-0132 | Hand-washing Policy |
| LAMC-PPP-0134 | Storage and Disposal of Chemical Hazardous Waste |
|  | Sure-Vue Color Mono Insert |

 |

|  |  |
| --- | --- |
| Author(s) | Alvin Castillo |