**ISED# 1 IS LIVE IN CERNER**

**9/28/2017**

**NOTES FOR NEW ESR INSTRUMENT**

* **TEST SITE: LA ESR iSED#1, all previous ESR orders that carries the old test site LA IrisTest#1 would have to be transferred to test site LA ESR iSED#1 using the Pending Inquiry mode in Cerner.**
* **Instrument results that is giving a result beyond linearity “less than 1” or “>130” in the print out will not cross to Cerner. Instead result will have to be manually entered as 0(zero) or 131, press enter and this will give a report of “<1” or “>130”in the result field. Click perform, then verify. DO NOT FREETEXT ANY RESULT. Free texting will inhibit abnormal flags in KPHC.**
* **PLEASE PERFORM iSED REPORTING NOTES in MTS.**