Radiometer ABL90 Flex Analyzer

Quality Control

##### PURPOSE	The purpose is to provide instructions on how to access, review and rerun ‘Built In Quality Control’ for the ABL90 Flex Analyzer.

##### _________________________________________________________________

##### POLICY	The ABL90 Flex Analyzer Built in Quality Control system must be analyzed every eight (8) hours of operation.  The internal Quality Control testing is scheduled as follows:

4:00 AM	S9030		Slot A

12:00 PM	S9040		Slot B

8:00 PM	S9050		Slot C

##### ______________________________________________________________________

##### SCOPE	The ABL90 Flex Analayzer is equipped with a quality management system, and all necessary checks are done automatically and continuously by the analyzer.  The Quality Management System is based on the following checks:

- Built in QC

- System Checks

- Analysis Check (performed before, under and after each activity)

Quality management is performed on six different solutions that are built into the solution pack.  Three solutions (CAL1, CAL2, CAL3) are used for calibration, analysis check and system check, and three solution (QC1 [S9030], QC2 [S9040], and QC3 [s9050] are used for quality controls.

Workplace	All laboratory employees are expected to maintain a safe working environment

Safety		and an injury-free workplace.  Laboratory employees are responsible for their

own safety, and the safety of others and adhering to all departmental and medical

center safety policies and procedures.

- For standard precautions and safety practices in the laboratory;  see LAMC-PPP-0123, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LAMC-PPP-0128, specifically, but not limited to, exposure to body fluids.

- For proper hand washing, see LAMC-PPP-0132, specifically, not limited to, proper hand washing.

- For proper infection control, see LAMC-PPP-0127, specifically, but not limited to, proper use of gloves.

Radiometer ABL90 Flex Analyzer

Quality Control

Continued..

_________________________________________________________________

- For proper handling of regular and infectious waste, see LAMC-PPP-0129, specifically, but not limited to, proper disposal of regular and bio hazardous waste.

- For proper cleaning of work area, see LAMC-PPP-0130 – Cleaning Work Are

- For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- For proper storage and disposal of chemical hazardous waste, see LAMC-PPP-0134.

REVIEW OF

QC RESULTS	Follow the steps below to access, review and rerun Built in QC results on

the ABL90 Flex Analyzer:

| Step | Action |
| --- | --- |
| 1 | Press ‘Menu’ on the screen. |
| 2 | Select Data Logs. |
| 3 | Select QC Logs. Each QC run will display with the date and time. |
| 4 | Select the desired run by highlighting the line. |
| 5 | Select ‘Result’ at the bottom of the display screen. Built in QC results will display. If any of the automatic QC run fails, select the ‘Stop Light’ icon. Highlight the failed QC and press ‘Start QC’ at the bottom to rerun automatic QC. |
| 6 | Select ‘Print’ if all results are acceptable and affixed the printout in the QC log. |
| 7 | Document corrective action taken for any QC failures in the ‘QC Action Log’ form. |

__________________________________________________________________

Radiometer ABL90 Flex Analyzer

Quality Control

Continued..

Controlled	The controlled documents are as follows:

Documents

| Document No. | Name of Documents |
| --- | --- |
| LAMC-PPP-0026 | Quality Manual |
| LAMC-PPP-0123 | Safety Practices |
| LAMC-PPP-0127 | Infection Control |
| LAMC-PPP-0128 | Universal Body Substance Precautions |
| LAMC-PPP-0129 | Handling of Regular and Infectious Waste |
| LAMC-PPP-0130 | Cleaning Work Areas |
| LAMC-PPP-0132 | Hand washing Policy |
| LAMC-PPP-0134 | Storage and disposal of Chemical Hazardous Waste |

##### NON			The Non Controlled Documents are as follows:

CONTROLLED

DOCUMENTS

| Document No. | Name of Documents |
| --- | --- |
|  | Radiometer ABL90 Flex Analyzer Operator’s Manual |

__________________________________________________________________

Author		Marianne Muyrong, CLS

Rosalie Fajardo, MS, CLS