#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer

| ##### Purpose |  |
| --- | --- |

- The purpose of this procedure is to describe preparation and staining sample for 10-color Lymphoma phenotyping using Navios Flow Cytometer.

Work Place Safety:

All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures.

- For standard precautions and safety practices in the laboratory; see LAMC-PPP-0123, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For proper handwashing, see LAMC-PPP-0132, specifically, but not limited to, proper handwashing.

- For proper handling of regular and infectious waste, see LAMC-PPP-0130, specifically, but not limited to proper disposal of regular and biohazardous waste.

- For proper cleaning of work area, see LAMC-PPP-0130- Cleaning Work Areas.

- For proper handling of chemicals and reagents, see the Chemical Hygiene Plan.

- For proper storage and disposal of chemical hazardous waste, see LAMC-PPP-0134.

| ##### Policy |  |
| --- | --- |

- Always wear personal protective equipment when working with Ammonium Chloride Lyse. This may include, but not limited to, protective eyewear, gloves and suitable laboratory attire when working with Ammonium Chloride.

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued.

| ##### Principle |  |
| --- | --- |

- This procedure describes the steps in performing cell surface marker study (immunophenotyping) performed from whole blood, bone marrow aspirate, effusion, solid tissue cell suspension, or other body fluids using panel of monoclonal antibodies conjugated with fluorochromes such Fluorescein Isothiocyanate (FITC), Phycoerythrin, (PE), Phycoerythrin-Texas Red® X (ECD), Phycoerythrin-Cy5 (PC5), PC-7, APC, APC Alexa-700, APC-Alexa-750, Pacific Blue, and Krome Orange.

- Immunophenotyping of leukocytes by flow cytometry requires removal of red blood cells, and this can be achieved by lysing erythrocytes. Ammonium chloride solution has been used for many years for lysis of red blood cells.

- Viable cell suspension is incubated with specific monoclonal antibodies, and antigens on the surface of the cell will react with the presence of their respective antibodies.

- After lysis of red cells, samples are stained with monoclonal antibodies and analyzed using flow cytometer by light scatter patterns, fluorescence intensity, and results are expressed as percentage of positive events in relation to all events acquired by electronic gating.

- The Navios is an instrument for phenotyping surface markers of white blood cells. The scatter plots of cells are measured and analyzed as cells pass through a LASER beam in a single-fine using the method of hydrodynamic focusing.

| ##### Specimen Requirement |  |
| --- | --- |

- Use cell suspension or body fluid less than 48 hours old.

- K3EDTA blood – 5ml. ACD and heparin is acceptable if EDTA anti-coagulated specimen is not available.

- Sodium heparinized bone marrow aspirate – minimum of 2ml. Wash once with Modified PBS before processing.

- Effusion (pericardial, pleural, or peritoneal) treated with anticoagulant with minimum volume of 10 ml.

- Body fluids cell suspension of washed or mashed solid tissue in cell media, specimen volume 5 to 10ml.

- CSF must be collected on CSF vial submit all available specimen, minimum volume 2ml.

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued

- Obtain WBC count, if available.

- Maintain and transport specimen at room temperature (18-22 0C) in a leak-proof container. Avoid temperature <10 0C and >37 0C.

- Take note on the worksheet and report on the form if the following occur:

- If the specimen is hot or cold but not obviously hemolyzed or frozen.

- If the specimen is hemolyzed or frozen.

- If clots are visible.

- If the specimen is >48 hours old at the time it arrives at the laboratory.

- Do viability testing. Refer to SOP LAMC-PPP-057001.

- Reagent, Materials and Equipment

| ##### Reagent: |  |
| --- | --- |

- Ammonium Chloride 500g Sigma-Aldrich Catalog number: 213330-500g.

- Potassium bicarbonate 327205 Sigma-Aldrich Catalog Number: 237205-500g

- Ethylenediaminetetraacetic acid tetrasodium salt.

- Orion Buffer (USA)

  - pH buffer, Orion 910104, Thermofisher

  - pH 7.00 buffer, Orion 910107, Thermofisher

  - pH10.01 buffer, Orion 910110, Thermofisher

  - pH Electrode Storage Solution, Orion 910001, Thermofisher

  - Distilled water

- Deionized Water

- Bleach

- IsoFlow Sheath Fluid, Beckman Coulter, PN 8547008

- Coulter Clenz cleaning agent, PN 8546930

- Bovine Serum Albumin 22% (BSA) Sigma-Adrich A7034-100ML

- Flow Set Pro Fluorospheres – Beckman Coulter PN A63492

- Phosphate buffered saline, pH =7.4. Sigma Catalog # P-3813

- CD Chex Plus Streck 213340

- Preparation:

| Step | Action |
| --- | --- |
| 1 | Take 1 packet (foil pouch) and dissolve in 1 liter of deionized water. Yield: 0.01 M PBS pH= 7.4 |
| 2 | Store in a 1 liter amber dispenser at room temperature. |

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued

- Monoclonal Antibody Reagents: Table 1 (Titered volume)

| Reagent | Catalog # | Manufacturer | ul/test |
| --- | --- | --- | --- |
| Kappa-FITC | A64828 | Beckman Coulter | 5 |
| Lambda-PE | A64827 | Beckman Coulter | 5 |
| CD19-ECD | IM2708U | Beckman Coulter | 5 |
| CD8-PC5.5 | A99019 | Beckman Coulter | 5 |
| CD5-PC7 | A99019 | Beckman Coulter | 2.5 |
| CD10-APC | IM3633 | Beckman Coulter | 2.5 |
| CD4-APC-AF700 | B10824 | Beckman Coulter | 5 |
| CD3-APC-AF750 | A66329 | Beckman Coulter | 5 |
| CD20-PB | A86355 | Beckman Coulter | 5 |
| CD45-KrO | A96416 | Beckman Coulter | 2.5 |
| CD57-FITC | IM466U | Beckman Coulter | 2.5 |
| CD16-PE | IM1238U | Beckman Coulter | 5 |
| CD56-PC5.5 | A79388 | Beckman Coulter | 5 |
| CD7-APC | A97050 | Beckman Coulter | 2.5 |
| CD8-PB | A82791 | Beckman Coulter | 5 |
| CD38-PC5 | A70205 | Beckman Coulter | 5 |
| CD23-APC-AF700 | A70207 | Beckman Coulter | 5 |
| CD20-APC-AF750 | A66331 | Beckman Coulter | 2.5 |
| FMC7-PB | B08466 | Beckman Coulter | 10 |
| Krome Orange | A96416 | Beckman Coulter | 2.5 |

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued

- Materials

- Spatula

- Weight boat

- Curwood Parafilm 4in W x125L, Catalog #13-374-10 Fisher-Scientific

- Parafilm Dispense/Cutter, Catalog# 06-666-52 Fisher Scientific

- Erlenmeyer Flask -1000ml

- 30 ml graduated beaker (3)

- 200 ml beaker

- Polypropylene round bottom tubes 12X75mm., Fisher Scientific, Catalog # 14-959-5

- Disposable plastic centrifuge, 15ml., Fisher Scientific, Catalog # 35-2099

- Disposable Petri dishes, size 100 x 15mm., Fisher Brand Catalog # 08-757-12

- Pipette Eppendorf

- Pipet tips (Finntip ® 250 Universal) Catalog # 9400260

- Transfer pipette, 1ml Graduated, Fisher Scientific, Catalog # 13-711-9A

- Test tube racks.

- Equipment

- pH Meter Fisher Sci Model 501

- Mettler XS Analytical Balance

- Navios Flow Cytometer

- Beckman Centrifuge

- Vortex Mixer

- Biosafety Cabinet

- Non-CO2  Incubator at 370C

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued

| ##### Quality Control |  |
| --- | --- |

See associated Standard Operating Procedures:

- Daily Start-up Procedure for Navios Flow Cytometer, LAMC-PPP-0535

- Quality Control Procedure for Navios Flow Cytometer, LAMC-PPP-0537

- Shutdown Procedure for Navios Flow Cytometer, LAMC-PPP-0536

- Navios Flow Cytometer, Preventive Maintenance, LAMC-PPP-0541

- Daily QC:

- Preparation of Flow Check Pro for Optical Alignment:

| - Step | - Action |
| --- | --- |
| - 1 | - Label a polypropylene round bottom tubes 12X75mm as Flow Check Pro indicate Lot number and expiration date on the tube. |
| - 2 | - Mix Flow Check Pro beads thoroughly, and add about 1ml of Flow Check Pro beads in the labeled Flow Check Pro tubes. |
| - 3 | - Load the labeled tube in the first position of the carousel. |

- Preparation of Flow Check Pro for Standardization of voltages of PMT.

| - Step | - Action |
| --- | --- |
| - 1 | - Label a polypropylene round bottom tubes 12X75mm as Flow Set Pro and indicate Lot number and expiration date on the tube. |
| - 2 | - Mix Flow Set Pro beads thoroughly, and add about 1ml of Flow Check Pro beads in the labeled Flow Check Pro tubes. |
| - 3 | - Load the labeled tube in the 2nd position of the carousel. |

- Preparation of Verify Tube using normal peripheral blood:

| - Step | - Action |
| --- | --- |
| - 1 | - Remove a vial of CD Chex Plus from the refrigerator and warm to room temperature (180 C to 300C) for 15 minutes before use. Mix vial between palms roll vial back and forth, hold vial between ends mix gently before sampling. |
| - 2 | - Stain 100ul of CD Chex Plus with using T- cell monoclonal antibody reagents. - VerifyTube:CD57-FITC/CD16-PE/CD5-PC7/CD7-APC/CD4-A700/CD3-A750/CD8-PB/CD45KO |
| - 3 | - Load the labeled tube in the 3rd position of the carousel. |

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued.

- How Acquire Data Optical Alignment, Standardization (Voltage and Mode), and compensation check.

| - Step | - Action |
| --- | --- |
| - 1 | - Load tube with 1ml of Flow Check Pro, Flow Set Pro, and Verify Tube in to a carousel. |
| - 2 | - Drag the 10-color daily QC Panel into the Acquisition Manager |
| - 3 | - Label Daily 10-color QC the file name column of the Acquisition Manager. |
| - 4 | - Press F9 to run the 10-color Daily QC Panel. |

How to prepare 10-Color Auto-setup tubes: (As needed only)

| Step | Action |
| --- | --- |
| 1 | Reconstitute 1 vials of Cytocomp cells (Part# 6607023) and into a 1ml Cytocomp Buffer. Let the solution sit for 5 to 10 minutes |
| 2 | Label 12x72mm tubes in the following manner: |

| Reagent | Catalog # | Manufacturer | ul/test |
| --- | --- | --- | --- |
| Flow Set Pro Fluorospheres | A63492 | Beckman Coulter | NA |
| CD45-FITC | IM0782U | Beckman Coulter | 5 |
| CD45-PE | IM2078U | Beckman Coulter | 5 |
| CD45-ECD | IM2710U | Beckman Coulter | 5 |
| CD45-PC5.5 | A54139 | Beckman Coulter | 5 |
| CD45-PC7 | IM3548U | Beckman Coulter | 5 |
| CD45-APC | IM2473U | Beckman Coulter | 5 |
| CD45-APC-AF700 | A71117 | Beckman Coulter | 5 |
| CD45-APC-AF750 | A71119 | Beckman Coulter | 5 |
| CD45-Pacific Blue | A74765 | Beckman Coulter | 5 |
| CD45-Krome Orange | A96416 | Beckman Coulter | 5 |
| VerifyTube:CD57-FITC/CD16-PE/CD5-PC7/CD7-APC/CD4-A700/CD3-A750/CD8-PB/CD45KO | See Table 1 Use Normal Peripheral blood for color compensation check. | Beckman Coulter | See Table 1 |

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued.

| 3 | Add corresponding volume CD45 antibody. |
| --- | --- |
| 4 | Incubate for 15minutes. |
| 5 | Add 2 ml pf PBS, spin for 1 minute using Clay Adams Centrifuge. |
| 6 | Decant supernatant and a 1 ml of PBS. |
| 7 | Run Auto-setup in Navios. |

How to run 10-color Auto-Setup: (as needed)

| Step | Action |
| --- | --- |
| 1 | From the Main menu, click on Tools and choose Auto-Setup Scheduler. |
| 2 | From Auto-Setup Scheduler window, choose 10-color Auto-Setup (10c-C) |
| 3 | Load Carousel. Enter carousel number and click on Schedule, click Close. |
| 4 | At Acquisition Manager (BT File name column), enter Flow Set Pro 10-C and at the Name column, Enter date. |
| 5 | Press F9 to run Auto-setup. The instrument will acquire Auto-Setup tubes. |

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued

Quality Control:

Ammonium Chloride performance as lysing agent must be check before lysing

Patient samples. The amount of time it takes to lyse red cells is measured. The

| Step | Action |
| --- | --- |
| 1 | Pipet 2-3 ml of sample (normal peripheral blood) in to a 15 ml conical tube. |
| 2 | Add 10 ml of freshly prepared NH4Cl (1X) |
| 3 | Mix the specimen after adding the lysing agent,and start timing for 7 minutes. |
| 4 | Check for completeness of erythrocyte lysing, by holding a printed page, and when the print can be read through the tube, then the lysis is complete. |
| 5 | Document the completeness of erythrocyte lysing by using a “NH4Cl pH Reading and Completeness of Erythrocyte Lysing” form. |

lysis time of erythrocyte must be complete within 7 minutes after addition of

Ammonium Chloride.

How to make 1X Ammonium Chloride solution:

| Step | Action |
| --- | --- |
| 1 | Make 1:10 Dilution of Lysing solution (NH4Cl) Add 9 volumes of de-ionized water to a 10X volume of stock solution. Do not refrigerate. |
| 2 | The 1X NH4Cl working solution surplus must be discarded after expiration date. Use adjust pH to 7.4. |

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued.

Note:  For bone marrow specimen and peripheral blood: To ensure ideal flow rate, the WBC     count should be from 5,000 to 10,000 cells/cu mm.  WBC above 15,000 requires adjustment

of the cell concentration to an appropriate level.

| WBC COUNT | DILUTION (with PBS) |
| --- | --- |
| 10,000 up to 20,000 | 1:2 |
| 20,000 up to 30,000 | 1:3 |
| 30,000 up to 40,000 | 1:4 |
| 40,000 up to 50,000 | 1:5 |
| 100,000 up to 200,000 | 1:10 |
| Above 200,000 | - 1:20 |

Note: Specimen with low WBC, use table below to estimate number tube/panel to stained:

Estimated Cell Counts: 10-color  (Lymphoma panel)

|  | Events/60 seconds | Estimated WBC Count x (103) |
| --- | --- | --- |
| No Flow | <20 events | .002 |
| 1 Tube | 21-200 | 0.002-0.02 |
| 2 Tubes | 201-400 | 0.02-0.04 |
| 3 Tubes | 401-600 | 0.04-0.06 |

| ##### Procedure: |  |
| --- | --- |

Kaiser SCPMG Flow Cytometry Lab 10-color Lymphoma Panel:

| 10-color Lymphoma Panel | FITC | PE | ECD | PC5.5 | PC7 | APC | APC-AF700 | APC-AF750 | PB | KrO |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Combo | Kappa 5ul | Lambda 5ul | CD19 5ul | CD8 5ul | CD5 2.5ul | CD10 2.5ul | CD4 5ul | CD3 5ul | CD20 5ul | CD45 2.5ul |
| T-Cell | CD57 2.5ul | CD16 5ul | CD19 5ul | CD56 5ul | CD5 2.5ul | CD7 2.5ul | CD4 5ul | CD3 5ul | CD8 5ul | CD45 2.5ul |
| B-Cell | Kappa 5ul | Lambda 5ul | CD19 5ul | CD38 5ul | CD5 2.5ul | CD10 2.5ul | CD23 5ul | CD20 2.5ul | FMC7 10ul | CD45 2.5ul |

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued.

How stain 10-color Lymphoma Panel:

| Step | Action |
| --- | --- |
| 1 | Lyse 2-3 ml of specimen using 10ml of 1X NH4Cl (working solution) for maximum of 7 minutes. |
| 2 | After 7 minutes, centrifuge lysed specimen for 5 minutes at 1500rpm. |
| 3 | Decant and add 10 ml of 1X NH4Cl, and spin then re-suspend with 1 ml of PBS and add 3 drops of 22% Albumin. |
| 4 | Incubate suspension of cells in the water bath (37% Celsius) for 20 minutes. |
| 5 | Wash cells with PBS 5 minutes at 1500rpm, decant and add 1 ml of RPMI. |
| 6 | Filter lysed samples using 210um filter to remove large tissue aggregates. |
| 7 | Obtain initial count of WBC and viability: - Add 100ul and 0.5ml of 50ugm/ml of Propidium Iodide solution. - Incubate for 5 minutes. - Run viability using Navios Flow Cytometer - adjust cell concentration to 5 to 10 cells x 103 cells/cumm. |
| 8 | Label 3 round bottom tubes (12x75mm): Combo, T-CELL, and B-cell |
| 9 | Add titered amount of (10-color Mab Panel) into each tube. (see table above) |
| 10 | Add 100ul of sample from Step 8 into each tube. |
| 11 | Incubate for 20 minutes |
| 12 | Add 2ml of PBS or NH4Cl. |
| 13 | Centrifuge tubes using Serofuge Clay Adams 2002 for 2 minutes. |
| 14 | Decant and add 0.5ml of PBS. |
| 15 | Run samples using Navios Flow Cytometer. |

- How to acquire sample for 10-color lymphoma.

| Step | Action |
| --- | --- |
| 1 | Clear Worklist by clicking New Worklist icon. |
| 2 | At the Resource Explorer, Click Folder containing 10 color lymphoma screen panel (BT Folder). |
| 3 | Drag 10-color lymphoma panel into Acquisition Manager. |
| 4 | At Acquisition Manager, type file names of samples at (BT# column): See file name format below: Year-Sample type-case number- last name XX- BT-XXX-XXXX. |
| 5 | Enter Carousel number, load Carousel, and Press F9. |
| 6 | - Check verify tube: Inspect verify tube to ensure proper color compensation: For Quadrant 1 and 3 (X-mean), the difference should be <0.3, and Quadrant 3 and 4 (Y-mean), the difference should be <0.3. Use form F303. |

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued.

| ##### Reporting and Documentation |  |
| --- | --- |

- LAMC-FORM-0146 for recording complete lysis of erythrocytes (NH4Cl).

- LAMC-Form-0102: LASER Usage Log Navios Flow Cytometer.

- LAMC-Form-0103: Daily Maintenance and QC.

- LAMC-Form-0104: Flow Check Pro

- Form F302: Flow Set Pro Log 10-color: Gains, Mode, and voltages of Navios.

- Form F303: Daily 10-color Lymphoma QC.

- Inspect verify tube to ensure proper color compensation: For Quadrant 1 and 3 (X-mean), the difference should be <0.3, and Quadrant 3 and 4 (Y-mean), the difference should be <0.3.

Controlled Document:

- LAMC-PPP-0123 - Safety Practices.

- LAMC-PPP-0132 – Handwashing Policy.

- LAMC-PPP-0127 – Infection Control.

- LAMC-PPP-0130 – Cleaning Work Areas.

- LAMC-FORM-0146 for recording complete lysis of erythrocytes (NH4Cl).

- LAMC-Form-0102: LASER Usage Log Navios Flow Cytometer.

- LAMC-Form-0103: Daily Maintenance and QC.

- LAMC-Form-0104: Flow Check Pro

- Form F302: Flow Set Pro Log 10-color: Gains, Mode, and voltages of Navios.

- Form F303: Daily 10-color Lymphoma QC.

| ##### Uncontrolled Document |  |
| --- | --- |

- RBC Lysing Solution and Cell Lysing Procedure: Retrieved, April 21, 2015: https://depts.washington.edu/flowlab/Cell%20Analysis%20Facility/RBC%20Lysing%20Solutions%20and%20Cell%20Lysing%20Procedure.pdf

- Lysing Solution IO Test 3 10X concentrate Ref A07799, January 01, 2003, Beckman Coulter.

#### 10-Color Lymphoma Immunophenotyping Using Navios Flow Cytometer, continued.

| ##### Self Assessment |  |
| --- | --- |

- What is maximum the lysing time of specimen with 1X ammonium chloride?

- 5 minutes

- 7 minutes

- It does not matter.

- To ensure proper color compensation:

- For Quadrant 1 and 3 (X-mean), the difference should be ≤0.3, and Quadrant 3 and 4 (Y-mean), the difference should be ≤0.3.

- For Quadrant 1 and 3 (X-mean), the difference should be <0.5, and Quadrant 3 and 4 (Y-mean), the difference should be <0.5.

- Color compensation check is not necessary, the instrument perform color compensation automatically.

- True or False:

WBC above 10,000 requires adjustment of the cell concentration to an appropriate level.

Author:

Giovanni B. Garcia, MBA, MT(ASCP), CLS