#### OPERATION OF pH Meter Fisher Sci Model 510

| ##### Purpose: |  |
| --- | --- |

- The purpose of this procedure is to describe the operation of pH meter Fisher Sci 510 for determining and adjusting pH of a lysing agent Ammonium Chloride NH4Cl.

Work Place Safety:

All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures.

- For standard precautions and safety practices in the laboratory; see LAMC-PPP-0123, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LAMC-PPP-0128, specifically, but not limited to, exposure to body fluids.

handwashing.

- For proper infection control, see LAMC-PPP-0127, specifically, but not limited to, proper use of gloves.

- For proper handling of regular and infectious waste, see LAMC-PPP-0129, specifically, but not limited to proper disposal of regular and biohazardous waste.

- For proper cleaning of work area, see LAMC-PPP-0130 - Cleaning Work Areas.

| ##### Policy: |  |
| --- | --- |

- Wear personal protective equipment when working with pH meter Fisher Sci 510

- pH Meter will be taken out of service when it is out of order.

- pH Meter has a one year warranty. Manufacturer For service call Fisher Sci at 310-743-9954.

#### OPERATION OF pH Meter Fisher Sci Model 510, continued

| ##### Principle: |  |
| --- | --- |

- pH is the negative log of the activity of the hydrogen ion concentration in aqueous solution Aqueous solutions with a pH less than 7 are said to be acidic and solutions with a pH greater than 7 are basic, and pure water has a hydrogen ion concentration of 7.

- Primary pH standard values are determined by measuring the voltage difference between a hydrogen electrode and a standard electrode such as the silver chloride glass electrode. By definition pH is the negative logarithm of the dissolved hydrogen atom and it is commonly expressed as the measure of the hydronium ion concentration.

- pH 510 meter is a microprocessor-based instrument for measuring pH and is equipped with liquid crystal display (LCD) to display pH reading.

| ##### Specimen Requirement: |  |
| --- | --- |

- Aqueous solution.

| ##### Reagent, Materials and Equipment: |  |
| --- | --- |

- Reagent:

- Ammonium Chloride 500g Sigma-Aldrich Catalog Number 213330-500g.

- Potassium bicarbonate 327205 Sigma-Aldrich Catalog Number 237205-500g

- Ethylenediaminetetraacetic acid tetrasodium salt.

- Orion Buffer

- pH 4.01 Buffer, Orion 910104, Thermofisher.

- pH 7.00 Buffer, Orion 910107, Thermofisher.

- pH 10.01 Buffer, Orion 910110, Thermofisher.

- pH Electrode Storage Solution, Orion 910001, Thermofisher

- Distilled Water

- Materials:

- Spatula

- Curwood Parafilm M 4in W x 125 L, Catalog #13-374-10 Fisher-Scientific.

- Parafilm Dispense/Cutter, Catalog # 06-666-52, Fisher Scientific.

- Erlenmeyer Flask -1000ul

- 25ml pyrex graduated beaker (3).

#### OPERATION OF pH Meter Fisher Sci Model 510, continued

- 100ml Pyrex graduated beaker.

- Equipment:

- pH Meter Fisher Sci Model 501

| ##### Procedure: |  |
| --- | --- |

How to perform three-point calibration of pH Meter Fisher Sci Model 510

| Step | Action |
| --- | --- |
| 1 | Turn on the power: Press the on/off key located at the front right side of the pH meter. |
| 2 | Rinse the pH electrode with deionized water or rinse solution. Note: Do not wipe the electrode as wiping may cause build-up of electrostatic charge and will create calibration and measurement instability. |
| 3 | Select first pH buffer, pH 7.0, and pour about 10ml of the buffer into a 25ml glass beaker. |
| 4 | Dip the electrode into the pH 7.0 buffer, and ensure that the electrode is completely immersed into the buffer; stir solution to create a homogenous solution. |
| 5 | Press CAL/MEAS key to enter pH Calibration mode. The CAL annunciator appears on top of the LCD. Note: The meter automatically recognizes the buffer standard set from SETUP mode. For this facility USA standard buffer is used: pH4.01, pH 7.00, and pH 10.01. |
| 6 | The stabilized pH measured value appears when the READY indicator displays on the left-hand corner. |
| 7 | Press ENTER key to confirm the calibration. The upper display flashes the calibration value for few seconds before storing the calibration point in the meter. |

#### OPERATION OF pH Meter Fisher Sci Model 510, continued

| 8 | Remove the electrode from the pH 7.0 buffer, rinse the electrode using distilled water to avoid contamination. |
| --- | --- |
| 9 | Place the electrode into the beaker containing pH 4.01. (The lower display of the meter automatically switches to the pH buffer solution selected. |
| 10 | The stabilized pH measured value (pH 4.01buffer) appears when the READY indicator displays on the left-hand corner. |
| 11 | Press ENTER key to confirm the calibration. The upper display flashes the calibration value for few seconds before storing the calibration point in the meter. |
| 12 | To perform the final buffer in the three-point calibration, remove the electrode from the pH 4.01 Buffer, rinse electrode and place and immerse electrode in pH 10.01 calibration buffer. |
| 13 | The stabilized pH measured value (pH 10.01buffer) appears when the READY indicator displays on the left-hand corner. |
| 14 | Press ENTER key to confirm the calibration. The upper display flashes the calibration value for few seconds before storing the calibration point in the meter. At this point, the meter is now ready to measure pH of an aqueous solution. The meter automatically returns to the pH measurement mode after three-point calibration is performed. |
| 15 | Record calibration reading on LAMC-FORM-0145. |

Note: If any error message appears during any of the calibration three-point, refer to section on “Troubleshooting pH meter calibration error.”

Windows of Accepted Values of USA calibration buffer:

| Calibration Buffer | Limit in (pH) |
| --- | --- |
| USA Calibration buffer (pH 4.01) | 4.0 ±1.0 |
| USA Calibration buffer (pH 7.01) | 7.0 ±1.5 |
| USA Calibration buffer (pH 10.01) | 10.0 ±1.0 |

#### OPERATION OF pH Meter Fisher Sci Model 510, continued

How to measure the pH of a solution:

| Step | Action |
| --- | --- |
| 1 | Rinse the electrode with distilled water or tap water before use to remove any impurities adhering to the surface of the electrode body. If the pH electrode has not been in the electrode solution (dehydrated), soak the electrode for half an hour into the electrode storage solution. |
| 2 | Press the ON key to switch the meter. The MEAS annunciator appears on top of the center of LCD. |
| 3 | Dip the electrode into the beaker or flask containing the sample, ensuring that the glass bulb of the electrode is completely immersed into the solution. Stir if necessary to create a homogenous sample. |
| 4 | Wait until the readings stabilize, and when the measurement stabilizes, the instrument displays a READY annunciator. At this point, take the pH reading. Note: The READ mode shows that the readings are stable within a range of pH 7.4 ±0.1 pH. - pH is not within range (pH 7.4 ±0.1), adjust pH as follows: if pH < 7.3 adjust pH using 1N NaOH, if pH is >7.5, adjust pH using 1N HCl. - Record pH readings using LAMC-FORM-0145. |

How to enter to SET-UP to customize pH meter preferences and defaults

| Step | Action |
| --- | --- |
| 1 | Turn off meter |
| 2 | With meter off, press and hold the MODE key, and press and release the ON key. The display will show SETUP indicator. Press up or down key to scroll through the program until you view P3.0 (low display and buffer set (buF)on the upper display. Press ENTER. |
| 3 | After pressing ENTER, the lower display will read either USA or nSt. |
| 4 | Press up or down scroll key to select buffer set (USA or nSt). |
| 5 | To confirm buffer standard, press ENTER. All characters on the display will flash momentarily and return to P3.0 |
| 6 | To exit P3.0 program press up down scroll key or press CAL/MEAS to return to measurement mode. |

#### OPERATION OF pH Meter Fisher Sci Model 510, continued

Troubleshooting pH meter calibration error.

Table below show a diagnosis guideline for possible problems encountered during calibration of pH meter.

| ERROR MESSAGE | INDICATES | CAUSE | CORRECTIVE ACTION |
| --- | --- | --- | --- |
| Err. 2 (in upper display) during Ion calibration | Too few calibration points | Meter calibrated with only 1-point in ion calibration | Recalibrate using 2 or more points |
| Err. 3 (in upper display) during Ion calibration | Slope error | Ion calibration solution not within acceptable range (15mV to 90mV, decade} | Recalibrate using fresh ion solutions Replace electrode. |
| Err. 4 (in upper display) during Ion calibration | Calibration points more than 1 decade apart | Calibration points more than 1decade apart | Recalibrate using calibration points 1 decade apart. |
| Err (annunciator) | Wrong keypad input | Button does not work in the current operation mode | Release key. Select valid buttons depending on mode. |
| Electrode icon (indicator) | Calibration error | Buffer value does not match value displayed or electrode is disconnected or failing | Use fresh buffer solution or check electrode connection . Possibly replace electrodes . |
| Ur or Or (in upper display) | Measured value is out of range | Electrode not connected Electrode clogged, dirty or broken Meter not calibrated . Measurement is under range or over range | Make sure electrode is connected Clean or replace electrode Recalibrate the meter. Treat samples to bring within meter measuring range |

#### OPERATION OF pH Meter Fisher Sci Model 510, continued

Other problems with pH meter:

| PROBLEM | PROBABLE CAUSE | REMEDIAL ACTION |
| --- | --- | --- |
| No display | - Mains power not switched on. - AC Adapter socket not inserted properly. | - Switch on the power supply. - Re-insert AC Adapter socket. |
| Unstable reading | a) Insufficient reference electrolyte in electrode. | a) Fill electrode with reference electrolyte. |
|  | b) Broken electrode | b) Replace electrode. |
|  | c) External 'noises' or induction ( e.g. electrical 'noise' caused by a nearby motor) | c) Remove or switch off interfering device. |
|  | d) Dirty electrode. | d) Clean the electrode. Hydrate, then rejuvenate if necessary. |
| Slow response | a) Dirty electrode. | a) Clean electrode. Hydrate, then rejuvenate if necessary. |
| Meter not responding to key press | - HOLD mode in operation. - Internal program error. | - Cancel HOLD mode. Press Hold button. - Reset all internal programs by removing and reinserting the A/C adapter. |

| ##### Interpretation |  |
| --- | --- |

- Not applicable for this procedure

| ##### Reporting and Documentation |  |
| --- | --- |

- After performing required daily maintenance, Daily Maintenance Form is signed daily by CLS and reviewed monthly by Laboratory Supervisor. (Maintenance Form is provided by Regional Instrumentation Laboratory.

| ##### Controlled Document |  |
| --- | --- |

- Daily Maintenance form provided by Kaiser Permanente SCMPG Laboratory Instrumentation Services.

- LAMC-PPP-0123 - Safety Practices.

- LAMC-PPP-0128 - Universal Body Substance Precautions (UBSP).

- LAMC-PPP-0132 -  Handwashing Policy.

#### OPERATION OF pH Meter Fisher Sci Model 510, continued

- LAMC-PPP-0127: Infection Control

- LAMC-PPP-0129: Handling Regular and Infections Waste.

- LAMC-PPP-0130: Cleaning Work Areas.

- LAMC-FORM-0145: Flow Cytometry pH meter Log Book (Fisher Sci 510)

| ##### Uncontrolled Document |  |
| --- | --- |

- Instruction Manual of pH Meter Fisher Sci Model 510. Retrieved April 2015 from http://www.nist.gov/ncnr/upload/pHMeter_FisherScientific_510.pdf.

| ##### Self-Assessment |  |
| --- | --- |

- Calibration of pH Meter Fisher 510 involves:

- One-point calibration

- Two-point calibration

- Three-point calibration

- The type of buffer used for calibration in the Flow Cytometry lab is:

- USA

- NIST

- USA and NIST

- Does not matter if USA or NIST

- True or False:

During the three-point calibration process, the operator press ENTER to confirm the calibration.

- READY annunciator display means:

- Measured pH value has been stabilized.

- Measured pH value needs to be stabilized by pressing Enter Key.

- All of the above.

Author: Giovanni B. Garcia, MT (ASCP), CLS, MBA