#### OPERATION OF METTLER XS204 ANALYTICAL BALANCE

| ##### Purpose: |  |
| --- | --- |

- This procedure describes steps involved in the operation of Mettler XS204 Analytical Balance.

| ##### Policy |  |
| --- | --- |

- Wear appropriate personal protective equipment when working with Mettler XS204 Flow Cytometer.

________________________________________________________________

Work Place Safety:

All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures.

- For standard precautions and safety practices in the laboratory; see LAMC-PPP-0123 specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For Universal Body Substance precautions, see LAMC-PPP-0128 specifically, but not limited to, exposure to body fluids.

- For proper handwashing, see LAMC-PPP-0132 specifically, but not limited to, proper handwashing.

- For proper infection control, see LAMC-PPP-0127 specifically, but not limited to, proper use of gloves.

- For proper handling of regular and infectious waste, see LAMC-PPP-0129 specifically, but not limited to proper disposal of regular and biohazardous waste.

- For proper cleaning of work area, see LAMC-PPP-0130 - Cleaning Work Areas.

| ##### Policy |  |
| --- | --- |

- Wear appropriate personal protective equipment when working with Navios Flow Cytometer.

| ##### Specimen Requirement |  |
| --- | --- |

- Not applicable for this procedure.

#### OPERATION OF METTLER XS204 ANALYTICAL BALANCE, continued

| ##### Reagent, Materials and Equipment |  |
| --- | --- |

- REAGENTS:

- Weights

- MATERIALS:

- Polypropylene tubes 12 X 75 mm.

- EQUIPMENT

- Mettler XS 204 Analytical Balance: Overview of the terminal:

- 1

- 6

- 2                                                                                       7

- 3                                                                                       3

- 5                4

- Home Key: This key takes the user directly to any menu to the active application.

- Directs the user to go to standard applications such as normal weighing, statistics, etc.

- Zero setting button: This button is used to set a new zero point.

- Tare key: This button is used to tare the balance.

- On/Off Key: Switches the balance on. Holding it down turns off the balance.

- Calls the menus for configuring active application.

- Transmits weighing results over the interface such as printer or PC.

#### OPERATION OF METTLER XS204 ANALYTICAL BALANCE, continued

| ##### Procedure: |  |
| --- | --- |

How to use Mettler XS204

| Step | Action |
| --- | --- |
| 1 | Press the On/Off key briefly to power on the unit. |
| 2 | Press Zero Display to set a new zero point. |
| 3 | Put a tare container on the balance and press “Tare” (See part #4) on the Overview Terminal then press the application menu button. |
| 4 | Open the glass draft shield on the side. |
| 5 | Place the object to be weighed on the weighing pan. |
| 6 | Close the glass draft shield. |
| 7 | As soon as the stability detector icon (small circle to the left of the weight display terminal) goes out, the display is stable, you may read the weighting result. |

| ##### Quality Control | Before weighing samples, perform weighing standards: 2g, 500mg, and 100mg to assess performance of the balance. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1 | Follow weighing procedure described on “How to use Mettler XS204” to weigh 2gram weight standard. Record results on the QC form for Mettler XS204. Use form LAMC-FORM-0061 |
| 2 | Follow weighing procedure described on “How to use Mettler XS204” to weigh 1gram weight standard. Record results on the QC form for Mettler XS204. Use from LAMC-FORM-0061 |
| 3 | Follow weighing procedure described on “How to use Mettler XS204” to weigh 100mg weight standard. Record results on the QC form for Metler XS204. Use form LAMC-FORM-0061 |

Note:  If balance fails, please check if the grid weighing pan is its proper place,

balance is centered, zero display, and tare is performed. Call for service

if balance continues to fail.

#### OPERATION OF METTLER XS204 ANALYTICAL BALANCE, continued

How to weigh samples in powder form:

| Step | Action |
| --- | --- |
| 1 | For installation remove the grid weighing pan from the weighing chamber. |
| 2 | Gently press the cover onto the grid weighing pan |
| 3 | Replace the grid weighing pan with installed grid weighing pan cover. |
| 4 | Put the specimen on a single-use weighing boat. |
| 5 | Press the On/Off key briefly to power on the unit. |
| 6 | Press Zero Display to set a new zero point. |
| 7 | Put a tare container on the balance and press “Tare” (See part #4) on the Overview Terminal then press the application menu button. |
| 8 | Open the glass draft shield on the side. |
| 9 | Place the weighing boat with sample to be weighed on the weighing pan. |
| 10 | Close the glass draft shield. |
| 11 | As soon as the stability detector icon (small circle to the left of the weight display terminal) goes out, the display is stable, you may read the weighting result. |

- How to weigh liquid object in a test tube:

| Step | Action |
| --- | --- |
| 1 | For installation remove the grid weighing pan from the weighing chamber. |
| 2 | Snap the ErgoClip on the grid weighing pan. |
| 3 | Replace the grid weighing pan with installed ErgoClip. |
| 4 | Press the On/Off key briefly to power on the unit. |
| 5 | Press Zero Display to set a new zero point. |
| 6 | Put a tare container on the balance and press “Tare” (See part #4) on the Overview Terminal then press the application menu button. |
| 7 | Open the glass draft shield on the side. |
| 8 | Place the object to be weighed on the ErgoClip. |
| 9 | Close the glass draft shield. |
| 10 | As soon as the stability detector icon (small circle to the left of the weight display terminal) goes out, the display is stable, you may read the weighting result. |

#### OPERATION OF METTLER XS204 ANALYTICAL BALANCE, continued

| ##### Interpretation |  |
| --- | --- |

- Not applicable for this procedure

| ##### Reporting & Documentation |  |
| --- | --- |

- Use Daily maintenance Maintenance and QC form: LAMC-FORM-0061

| ##### Controlled Document |  |
| --- | --- |

- LAMC-PPP-0123 LGM 8000: Safety Practices.

- LAMC-PPP-0132 LGM 8010: Handwashing Policy.

- LAMC-PPP-0127 LGM 8004: Infection Control

- LAMC-PPP-0129 LGM 8006: Handling Regular and Infections Waste.

- LAMC-PPP-0130 LGM 8007: Cleaning Work Areas.

- LAMC-PPP-0130: Cleaning Work Areas

- LAMC-FORM-0061: Mettler XS204 Quality Control Form

#### OPERATION OF METTLER XS204 ANALYTICAL BALANCE, continued

| ##### Uncontrolled Document |  |
| --- | --- |

| Document | Description |
| --- | --- |
| XS Model Part 1 | Excellence Analytical Balances XS Models Part 1 Mettler- Toledo AG, Laboratory and Weighing Technologies Greifensee, Switzerland, 05/2012 |
| XS Model Part 2 | Excellence Analytical Balances XS Models Part 2 Mettler- Toledo AG, Laboratory and Weighing Technologies Greifensee, Switzerland, 10/2010 |
| XS Model Part 3 | Excellence Analytical Balances XS Models Part 3 Mettler- Toledo AG, Laboratory and Weighing Technologies Greifensee, Switzerland, 10/2010 |

| ##### Self-Assessment |  |
| --- | --- |

- 1. To power on the analytical balance:

- To turn on the balance press On/Off key, and to turn if off press and hold On/Off key.

- To turn on the balance press and hold the the On/Off key, and to turn it off, press On/Off Key.

- a and b is correct.

- a is correct

- b is correct.

- True or False: As soon as the stability detector icon (small circle to the left of the weight display terminal) goes out, the display is stable, you may read the weighting result.

- Author:

- Giovanni B. Garcia, CLS, MT(ASCP), MBA