#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer

| ##### Purpose |  |
| --- | --- |

- The purpose of this procedure is to describe preparation and staining sample for 10-color Lymphoma phenotyping using Navios Flow Cytometer.

Work Place Safety:

All laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures.

- For standard precautions and safety practices in the laboratory; see LAMC-PPP-0123, specifically, but not limited to, equipment safety, proper body mechanics, sharps exposure and proper use of personal protective equipment (PPE).

- For proper handwashing, see LAMC-PPP-0132, specifically, but not limited to, proper handwashing.

- For proper handling of regular and infectious waste, see LAMC-PPP-0130, specifically, but not limited to proper disposal of regular and biohazardous waste.

- For proper cleaning of work area, see LAMC-PPP-0130- Cleaning Work Areas.

- For proper handling of chemicals and reagents and disposal of Hazardous Waste. see the Chemical Hygiene Plan.LAMC-PPP-0134.

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

| ##### Principle |  |
| --- | --- |

- This procedure describes the steps in performing cell surface marker study (immunophenotyping) performed from whole blood, bone marrow aspirate, effusion, solid tissue cell suspension, or other body fluids using panel of monoclonal antibodies conjugated with fluorochromes such Fluorescein Isothiocyanate (FITC), Phycoerythrin, (PE), Phycoerythrin-Texas Red® X (ECD), Phycoerythrin-Cy5 (PC5.5), PC-7, APC, APC Alexa-700, APC-Alexa-750, Pacific Blue, and Krome Orange.

- Immunophenotyping of leukocytes by flow cytometry requires removal of red blood cells, and this can be achieved by lysing erythrocytes. Ammonium chloride solution has been used for many years for lysis of red blood cells.

- Viable cell suspension is incubated with specific monoclonal antibodies, and antigens on the surface of the cell will react with the presence of their respective antibodies.

- After lysis of red cells, samples are stained with monoclonal antibodies and analyzed using flow cytometer by light scatter patterns, fluorescence intensity, and results are expressed as percentage of positive events in relation to all events acquired by electronic gating.

- The Navios is an instrument for phenotyping surface markers of white blood cells. The scatter plots of cells are measured and analyzed as cells pass through a LASER beam in a single-fine using the method of hydrodynamic focusing.

| ##### Specimen Requirement |  |
| --- | --- |

- Use cell suspension or body fluid less than 48 hours old.

- K3EDTA blood – 2ml minimum.  ACD and heparin is acceptable if EDTA anti-coagulated specimen is not available.

- Sodium heparinized bone marrow aspirate – minimum of 2ml. Wash once with Modified PBS before processing.

- Effusion (pericardial, pleural, or peritoneal) may be treated with anticoagulant with minimum volume of 10 ml.

- Body fluids cell suspension of washed or mashed solid tissue in cell media, specimen volume 5 to 10ml.

- CSF must be collected on CSF vial submit all available specimen, minimum volume 2ml.

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

- Obtain WBC count, if available.

- Maintain and transport specimen at room temperature (18-22 0C) in a leak-proof container. Avoid temperature <10 0C and >37 0C.

- Take note on the worksheet and report on the form if the following occur:

- If the specimen is hot or cold but not obviously hemolyzed or frozen.

- If the specimen is hemolyzed or frozen.

- If clots are visible.

- If the specimen is >48 hours old at the time it arrives at the laboratory.

- Reagent, Materials and Equipment

| ##### Reagent: |  |
| --- | --- |

- Ammonium Chloride 500g Sigma-Aldrich Catalog number: 213330-500g.

- Potassium bicarbonate 327205 Sigma-Aldrich Catalog Number: 237205-500g

- Ethylenediaminetetraacetic acid tetrasodium salt.

- Orion Buffer (USA)

  - pH 4.0 buffer, Orion 910104, Thermofisher

  - pH 7.00 buffer, Orion 910107, Thermofisher

  - pH10.01 buffer, Orion 910110, Thermofisher

  - pH Electrode Storage Solution, Orion 910001, Thermofisher

  - Distilled water

- Deionized Water

- Bleach

- IsoFlow Sheath Fluid, Beckman Coulter, PN 8547008

- Coulter Clenz cleaning agent, PN 8546930

- CRISP TdT Control – Phoenix Flow – P/N CRTDP-100

- Bovine Serum Albumin 22% (BSA) Sigma-Adrich A7034-100ML

- Flow Set Pro Fluorospheres – Beckman Coulter PN A63492

- Flow Check Fluorospheres – Beckman Coulter PN A63493

- CD-Chex Plus Streck 213340

- Phosphate buffered saline, pH =7.4. Sigma Catalog # P-3813

- Fix and Perm Medium A -Life Technologies GAS0015100

- Fix and Perm Medium B – Life Technologies GAS002510

- Preparation:

| Step | Action |
| --- | --- |
| 1 | Take 1 packet (foil pouch) and dissolve in 1 liter of deionized water. Yield: 0.01 M PBS pH= 7.4 |
| 2 | Store in a 1liter amber dispenser at room temperature. |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

- Monoclonal Antibody Reagents for Leukemia Panel: Table 1 (Titered volume)

| Panel | Reagent | Catalog # | Manufacturer | ul/test |
| --- | --- | --- | --- | --- |
| AML Panel | None | Not Applicable (NA) | NA | NA |
|  | None | NA | NA | NA |
|  | CD64-ECD | A98434 | Beckman Coulter | 1.0 |
|  | CD117-PC5.5 | A66333 | Beckman Coulter | 1.0 |
|  | CD33-PC7 | A54824 | Beckman Coulter | 1.0 |
|  | CD13-APC | A87783 | Beckman Coulter | 1.0 |
|  | CD34-APC-AF700 | A86354 | Beckman Coulter | 2.0 |
|  | CD14-APC-AF750 | A86052 | Beckman Coulter | 1.0 |
|  | HLA-DR-PB | A74781 | Beckman Coulter | 1.0 |
|  | CD45-KrO | A96416 | Beckman Coulter | 2.5 |
| ALL Panel | CD2-FITC | IM0442U | Beckman Coulter | 10 |
|  | CD34-PE | 340669 Clone 8G12 | Becton Dickinson | 10 |
|  | CD3-ECD | IM2705U | Beckman Coulter | 1.0 |
|  | CD5-PC5.5 | A70203 | Beckman Coulter | 1.0 |
|  | CD7-PC7 | A46526 | Beckman Coulter | 1.0 |
|  | CD22-APC | A60791 | Beckman Coulter | 1.0 |
|  | CD19-APC-AF700 | A78837 | Beckman Coulter | 1.0 |
|  | CD10-APC-AF750 | A89310 | Beckman Coulter | 5.0 |
|  | CD20-PB | A74777 | Beckman Coulter | 2.5 |
|  | CD45-KrO | A96416 | Beckman Coulter | 2.5 |
| Intracellular | cyTDT-FITC | F713950 | Agilent | 5.0 |
|  | NA | NA | NA | NA |
|  | cyCD22-ECD | B10245 | Beckman Coulter | 1.0 |
|  | cyMPO-PC5.5 | B23132 | Beckman Coulter | 1.0 |
|  | NA | NA | NA | NA |
|  | cyCD79a-APC | A60793 | Beckman Coulter | 1.0 |
|  | NA | NA | NA | NA |
|  | cyCD3-APC-AF750 | A66329 | Beckman Coulter | 1.0 |
|  | NA | NA | NA | NA |
|  | sCD45-KrO | A96416 | Beckman Coulter | 2.5 |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

- Materials

- Spatula

- Weight boat

- Curwood Parafilm 4in W x125L, Catalog #13-374-10 Fisher-Scientific

- Parafilm Dispense/Cutter, Catalog# 06-666-52 Fisher Scientific

- Erlenmeyer Flask -1000ml and 4000ml

- 30 ml graduated beaker (3)

- 200 ml beaker

- Polypropylene round bottom tubes 12X75mm., Fisher Scientific, Catalog # 14-959-5

- Disposable plastic centrifuge, 15ml., Fisher Scientific, Catalog # 35-2099

- Disposable Petri dishes, size 100 x 15mm., Fisher Brand Catalog # 08-757-12

- Pipette Eppendorf

- Pipet tips (Finntip ® 250 Universal) Catalog # 9400260

- Transfer pipette, 1ml Graduated, Fisher Scientific, Catalog # 13-711-9A

- Test tube racks.

- Equipment

- pH Meter Fisher Sci Model 510

- Mettler XS Analytical Balance

- Navios Flow Cytometer

- Beckman Centrifuge

- Vortex Mixer

- Biosafety Cabinet

- Non-CO2 Incubator at 370C

- Stirrer and stirrer bar

| ##### Quality Control |  |
| --- | --- |

See associated Standard Operating Procedures:

- Daily Start-up Procedure for Navios Flow Cytometer, LAMC-PPP-0535

- Shutdown Procedure for Navios Flow Cytometer, LAMC-PPP-0536

- Navios Flow Cytometer, Preventive Maintenance, LAMC-PPP-0541

- 10-color Lymphoma Immunophenotyping Using Navios Flow Cytometer LAMC-PPP-0792

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued,

- Daily QC:

- Preparation of Flow Check Pro for Optical Alignment:

| - Step | - Action |
| --- | --- |
| - 1 | - Label a polypropylene round bottom tubes 12X75mm as Flow Check Pro indicate Lot number and expiration date on the tube. |
| - 2 | - Mix Flow Check Pro beads thoroughly and add about 1ml of Flow Check Pro beads in the labeled Flow Check Pro tube. |
| - 3 | - Load the labeled tube in the first position of the carousel. |

- Preparation of Flow Check Pro for Standardization of voltages of PMT.

| - Step | - Action |
| --- | --- |
| - 1 | - Label a polypropylene round bottom tubes 12X75mm as Flow Set Pro and indicate Lot number and expiration date on the tube. |
| - 2 | - Mix Flow Set Pro beads thoroughly and add about 1ml of Flow Check Pro beads in the labeled Flow Check Pro tubes. |
| - 3 | - Load the labeled tube in the 2nd position of the carousel. |

- Preparation of Verify Tube using CD Chex Plus:

| - Step | - Action |
| --- | --- |
| - 1 | - Take CD Chex Plus from the storage refrigerator, bring to room temperature. Note: Open vial stability is 30 day. Record expiration on form LAMC-PPP-0149. |
| - 2 | - Stain CD Chex Plus with using T- cell monoclonal antibody reagents. - VerifyTube:CD57-FITC/CD16-PE/CD5-PC7/CD7-APC/CD4-A700/CD3-A750/CD8-PB/CD45KO |
| - 3 | - Load the labeled tube in the 3rd position of the carousel. |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

- How Acquire Data Optical Alignment, Standardization (Voltage and Mode), and compensation check.

| - Step | - Action |
| --- | --- |
| - 1 | - Load tube with 1ml of Flow Check Pro, Flow Set Pro, and Verify Tube into a carousel. |
| - 2 | - Drag the 10-color daily QC Panel into the Acquisition Manager |
| - 3 | - Label Daily 10-color QC the file name column of the Acquisition Manager. |
| - 4 | - Press F9 to run the 10-color Daily QC Panel. |

How to prepare 10-Color Auto-setup tubes: (As needed only)

| Step | Action |
| --- | --- |
| 1 | Reconstitute 1 vials of Cytocomp cells (Part# 6607023) and into a 1ml Cytocomp Buffer. Let the solution sit for 5 to 10 minutes |
| 2 | Label 12x72mm tubes in the following manner: |

| Reagent | Catalog # | Manufacturer | ul/test |
| --- | --- | --- | --- |
| Flow Set Pro Fluorospheres | A63492 | Beckman Coulter | NA |
| CD45-FITC | IM0782U | Beckman Coulter | 5 |
| CD45-PE | IM2078U | Beckman Coulter | 5 |
| CD45-ECD | IM2710U | Beckman Coulter | 5 |
| CD45-PC5.5 | A54139 | Beckman Coulter | 5 |
| CD45-PC7 | IM3548U | Beckman Coulter | 5 |
| CD45-APC | IM2473U | Beckman Coulter | 5 |
| CD45-APC-AF700 | A71117 | Beckman Coulter | 5 |
| CD45-APC-AF750 | A71119 | Beckman Coulter | 5 |
| CD45-Pacific Blue | A74765 | Beckman Coulter | 5 |
| CD45-Krome Orange | A96416 | Beckman Coulter | 5 |
| VerifyTube:CD57-FITC/CD16-PE/CD5-PC7/CD7-APC/CD4-A700/CD3-A750/CD8-PB/CD45KO | For cells use CD Chex Plus 213326 | Streck | See page 10 of LAMC-PPP-0792 |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

| 3 | Add corresponding volume CD45 antibody. |
| --- | --- |
| 4 | Incubate for 15minutes. |
| 5 | Add 2 ml pf PBS, spin for 1 minute using Clay Adams Centrifuge. |
| 6 | Decant supernatant and a 1 ml of PBS. |
| 7 | Run Auto-setup in Navios. |

How to run 10-color Auto-Setup: (as needed)

| Step | Action |
| --- | --- |
| 1 | From the Main menu, click on Tools and choose Auto-Setup Scheduler. |
| 2 | From Auto-Setup Scheduler window, choose 10-color Auto-Setup. |
| 3 | Load Carousel. Enter carousel number and click on Schedule, click Close. |
| 4 | At Acquisition Manager (BT File name column), enter Flow Set Pro 10-C and at the Name column, Enter date. |
| 5 | Press F9 to run Auto-setup. The instrument will acquire Auto-Setup tubes. |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

Quality Control:

Ammonium Chloride performance as lysing agent must be check before lysing

Patient samples. The amount of time it takes to lyse red cells is measured. The

| Step | Action |
| --- | --- |
| 1 | Pipet 2-3 ml of sample (normal peripheral blood) in to a 15 ml conical tube. |
| 2 | Add 10 ml of freshly prepared NH4Cl (1X) |
| 3 | Mix the specimen after adding the lysing agent and start timing for 7 minutes. |
| 4 | Check for completeness of erythrocyte lysing, by holding a printed page, and when the print can be read through the tube, then the lysis is complete. |
| 5 | Document the completeness of erythrocyte lysing by using a “NH4Cl pH Reading and Completeness of Erythrocyte Lysing” form. |

lysis time of erythrocyte must be complete within 7 minutes after addition of

Ammonium Chloride.

How to make 1X Ammonium Chloride solution:

| Step | Action |
| --- | --- |
| 1 | Make 1:10 Dilution of Lysing solution (NH4Cl) Add 9 volumes of de-ionized water to a 10X volume of stock solution. Do not refrigerate. |
| 2 | The 1X NH4Cl working solution surplus must be discarded after expiration date. Use adjust pH to 7.4± 0.1. |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

Note:  For bone marrow specimen and peripheral blood: To ensure ideal flow rate, the WBC     count should be from 5,000 to 10,000 cells/cu mm.  WBC above 15,000 requires adjustment

of the cell concentration to an appropriate level.

| WBC COUNT | DILUTION (with PBS) |
| --- | --- |
| 10,000 up to 20,000 | 1:2 |
| 20,000 up to 30,000 | 1:3 |
| 30,000 up to 40,000 | 1:4 |
| 40,000 up to 50,000 | 1:5 |
| 100,000 up to 200,000 | 1:10 |
| Above 200,000 | - 1:20 |

For specimens with low WBC use the Table below to determine number of panel to stained:

Note: Spin to concentrate cells and q.s. to 0.5ml.

| Number of Panels | Events/60 seconds | Approximate WBC X 104 |
| --- | --- | --- |
| No flow | <200 | 0.002 |
| 1 Panel/tube | 201-600 | 0.02-0.06 |
| 2 Panels/tubes | 601-1,000 | 0.06-.1 |
| 3 Panels/tubes | 1,000-2,000 | 0.1-0.2 |

Note: Concentrate cells before staining.

| ##### Procedure: |  |
| --- | --- |

Kaiser SCPMG Flow Cytometry Lab 10-color Myeloid and Lymphoblastic 10-color Panel:

| 10-color Leukemia Panel | FITC | PE | ECD | PC5.5 | PC7 | APC | APC-AF700 | APC-AF750 | PB | KrO |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| AML Panel | NONE | NONE | CD64 1.0ul | CD117 1.0ul | CD33 1.0ul | CD13 1.0ul | CD34 2.0ul | CD14 1.0ul | HLDR 1.0 | CD45 2.5ul |
| ALL Panel | CD2 10.0ul | CD34 10.0ul | CD3 1.0ul | CD5 1.0ul | CD7 1.0ul | CD22 1.0ul | CD19 1.0ul | CD10 5.0ul | CD20 2.5ul | CD45 2.5ul |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

How stain 10-color Myeloid and Lymphoblastic 10-color Panel:

| Step | Action |
| --- | --- |
| 1 | Lyse 2-3 ml of specimen using 10ml of 1X NH4Cl (working solution) for 7 minutes. |
| 2 | After 7 minutes, centrifuge lysed specimen for 5 minutes at 1500rpm. |
| 3 | Decant and add 10 ml of 1X NH4Cl, centrifuge for 5 minutes at 1500rpm. |
| 4 | Add 10 ml of PBS and vortex. |
| 5 | Wash cells with PBS 5 minutes at 1500rpm, decant and add 1 ml of RPMI. |
| 6 | Filter lysed samples using 210um filter to remove large tissue aggregates. |
| 7 | Obtain initial count of WBC and viability: - Add 100ul and 0.5ml of 50ugm/ml of Propidium Iodide solution. - Incubate for 5 minutes. - Run viability using Navios Flow Cytometer - adjust cell concentration to 5 to 10 cells x 103 cells/cumm. |
| 8 | Label round bottom tubes (12x75mm): Based from the flow cytometry test ordered by the hematopathologist: Myeloid and T-Lymphoblastic or both. |
| 9 | Add titered amount of (10-color Mab Panel) into each tube. (see table above) |
| 10 | Add 100ul of sample from Step 8 into each tube. |
| 11 | Incubate for 20 minutes |
| 12 | Add 2ml of PBS or 1X NH4Cl. |
| 13 | Centrifuge tubes using Serofuge Clay Adams 2002 for 2 minutes. |
| 14 | Decant and add 0.5ml of PBS. |
| 15 | Run samples using Navios Flow Cytometer. |

- How to acquire sample for 10-color Myeloid Panel and 10-color T-Lymphoblastic Panel.

| Step | Action |
| --- | --- |
| 1 | Clear Worklist by clicking New Worklist icon. |
| 2 | At the Resource Explorer, Click Folder containing 10 color Leukemia screen panel (BT Folder). |
| 3 | Drag 10-color appropriate 10-color panel into Acquisition Manager. Note: For 10-color AML Panel, use 10-color AML Panel. For 10-color ALL Panel, use 10-C ALL Panel. |
| 4 | At Acquisition Manager, type file names of samples at (BT# column): See file name format below: Year-Sample type-case number- last name XX- BT-XXXXX-XXXX. |
| 5 | Enter Carousel number, load Carousel, and Press F9. |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer, continued

How to stain Intracellular Panel:

| Panel | FITC | PE | ECD | PC5.5 | PC7 | APC | APC-AF700 | APC-AF750 | PB | KrO |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Intracellular | cyTdT 5.0ul | NONE | cyCD22 1.0ul | cyMPO 1.0ul | NONE | cyCD79a 1.0ul | NONE | cyCD3 1.0ul | NONE | sCD45 2.5ul |

| Step | Action |
| --- | --- |
| 1 | Lyse 2-3 ml of specimen using 10ml of 1X NH4Cl (working solution) for 7 minutes. |
| 2 | After 7 minutes, centrifuge lysed specimen for 5 minutes at 1500rpm. |
| 3 | Decant and add 10 ml of 1X NH4Cl, centrifuge for 5 minutes at 1500rpm. |
| 4 | Add 10 ml of PBS and vortex. |
| 5 | Wash cells with PBS 5 minutes at 1500rpm, decant and add 1 ml of RPMI. |
| 6 | Filter lysed samples using 210um filter to remove large tissue aggregates. |
| 7 | Obtain initial count of WBC and viability: - Add 100ul and 0.5ml of 50ugm/ml of Propidium Iodide solution. - Incubate for 5 minutes. - Run viability using Navios Flow Cytometer - adjust cell concentration to 5 to 10 cells x 103 cells/cumm. |
| 8 | Label round bottom tubes (12x75mm): Based from the flow cytometry test ordered by the hematopathologist: Intracellular Acute Leukemia and B-Lymphoblastic L or both. |
| 9 | Add titered amount for Surface Mab Panel into each tube. (see table above) Note: For Surface markers for Intracellular Acute: sCD45 (2.5 ul) |
| 10 | Add 100ul of sample from Step 9 into each tube and incubate for 15 minutes in the dark at room temperature. Wash cells with PBS 5 minutes at 1500rpm, decant and add 1 ml of RPMI. |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer

| 11 | For intracellular Stain: a) Add 100 ul of Reagent A (Fixation Medium) and incubate for 15 minutes at room temperature. b) Wash once with 2ml PBS, with centrifuge for 5 minutes at 1500 rpm decant supernatant, and vortex to fully resuspend the cell pellet. c) Add 100ul of Reagent B (Permeabilization Medium) and the recommended volume of conjugated intracellular antibody. See table of Intracellular Acute Leukemia Panel. Note: Intracellular markers for 10-color Acute Leukemia Panel: cyTdT, cyCD22, cyMPO, cyCD79a, and cyCD3. e) Vortex for 1 to 2 seconds and incubate for 20 minutes. f) Wash once in 2ml PBS. Centrifuge for 5 minutes at 1500 rpm decant supernatant, and resuspend cells in a sheath |
| --- | --- |

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer

- How to acquire sample for 10-color Intracellular Acute Leukemia Panel and 10-color B-Lymphoblastic Panel.

| Step | Action |
| --- | --- |
| 1 | Clear Worklist by clicking New Worklist icon. |
| 2 | At the Resource Explorer, Click Folder containing 10 color Leukemia screen panel (BT Folder). |
| 3 | Drag 10-color appropriate 10-color panel into Acquisition Manager. Note: For 10-color Intracellular Panel, use 10-C Intracellular Panel. |
| 4 | At Acquisition Manager, type file names of samples at (BT# column): See file name format below: Year-Sample type-case number- last name XX- BT-XXXXX-XXXX. |
| 5 | Enter Carousel number, load Carousel, and Press F9. |

| ##### Reporting and Documentation |  |
| --- | --- |

- LAMC-FORM-0145: for pH meter readings.

- LAMC-FORM- 0146: for recording complete lysis of erythrocytes.

- LAMC-FORM-0147: Flow Check Pro Optical Alignment beads.

- LAMC-FORM-0418:  Standardization use Flow Set Pro Flourospheres.

- LAMC-FORM-0149: Verification – CD-Chex Plus, 10-color QC

- LAMC-FORM-0150: Summary Shift/Trend Monitoring 10-color Flow Cytometry

- LAMC-FORM-0102: LASER Usage Log Navios Flow Cytometer.

- LAMC-FORM-0103: Daily Maintenance and QC.

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer

Controlled Document:

- LAMC-PPP-0123 - Safety Practices.

- LAMC-PPP-0132 – Handwashing Policy.

- LAMC-PPP-0127 – Infection Control.

- LAMC-PPP-0130 – Cleaning Work Areas.

- LAMC-FORM-0145 for pH meter readings.

- LAMC-FORM- 0146: for recording complete lysis of erythrocytes.

- LAMC-FORM-0147: for Flow Check Pro Optical Alignment beads.

- LAMC-FORM-0418: Standardization use Flow Set Pro Flourospheres.

- LAMC-FORM-0149:  Verification – CD-Chex Plus, 10-color QC

- LAMC-FORM-0150: Summary Shift/Trend Monitoring 10-color Flow Cytometry

- LAMC-FORM-0102: LASER Usage Log Navios Flow Cytometer.

- LAMC-FORM-0103: Daily Maintenance and QC

| ##### Uncontrolled Document |  |
| --- | --- |

- RBC Lysing Solution and Cell Lysing Procedure: Retrieved, April 21, 2015: https://depts.washington.edu/flowlab/Cell%20Analysis%20Facility/RBC%20Lysing%20Solutions%20and%20Cell%20Lysing%20Procedure.pdf

- Lysing Solution IO Test 3 10X concentrate Ref A07799, January 01, 2003, Beckman Coulter.

#### 10-Color Acute Leukemia Immunophenotyping Using Navios Flow Cytometer

| ##### Self- Assessment: |  |
| --- | --- |

- How long is the lysing time of specimen with 1X ammonium chloride?

- 10 minutes

- Maximum of 7 minutes

- It does not matter.

- True or False:

WBC above 15,000 requires adjustment of the cell concentration to an appropriate level.

- In intracellular staining, the incubation time after adding Reagent A (Fixation Medium):

- 10 minutes

- 15 minutes

- 20 minutes

- No incubation needed.

- In intracellular staining, the incubation time after adding Reagent B (Permeabilization Medium):

- 10 minutes

- 15 minutes

- 20 minutes

- No incubation needed.

Author:

Giovanni B. Garcia, MBA, MT(ASCP), CLS