**URINE FENTANYL QUICK BENCH GUIDE**

|  |  |
| --- | --- |
| **Reagent:** | IMMUNALYSIS Sefria Fentanyl Urine Ref. no. 618UR-002:Compartment A: **EA** (Enzyme Acceptor) & Compartment B: **ED** (Enzyme Donor) |
| **Onboard stability:** | 28 Days |
| **Calibrator:** | * Fentanyl Negative Calibrator REF# NEG-600-5
* Fentanyl Urine Calibrator 2 (2ng/mL) REF# 10020-5
* Unopened stability: Until printed expiration if stored at 2-8°C
* Opened stability: 60 days if stored at 2-8°C
 |
| **Calibration Frequency:** | * **At least every 7 days**
* If reagent is a new lot
* If QC is out of range or shows unusual trend or shift
 | * After major preventative maintenance or change of critical instrument component
 |
| **Controls:** | * **Fentanyl Urine Calibrator 1** (1ng/mL) REF# 10019-5 **(NEGATIVE QC)**
* ***(Result must be <2.0 to indicate instrument is correctly producing a NEGATIVE result)***
* **Fentanyl Urine Calibrator 3** (4ng/mL) REF# 10021-5 **(POSITIVE QC)**
* ***(Result must be >2.0 to indicate instrument is correctly producing a POSITIVE result)***
* Unopened stability: Until printed expiration if stored at 2-8°C
* Opened stability: 60 days if stored at 2-8°C
 |
| **Control Frequency:** | * **Every 12 hours included in SET B**
* After performing calibration
 | * To verify instrument operation
 |
| **Specimen type:** | * Urine – mix well in primary sterile container, then aliquot and centrifuge to remove gross particulate matter.
 |
| **Expected results:** | * If result is <2.0 ng/mL: **NEGATIVE**
* If result is >2.0 ng/mL: **POSITIVE**
 |
| **Result Interpretation by Cerner** |

|  |  |  |
| --- | --- | --- |
| If result is... | And… | Then report result as... |
| < 2.0  | Any patient | Negative |
| ≥ 2.0 | Patient is < 14 years old OR \*Patient is a new mother | Preliminary Positive. Pending Confirmation**Note:** OPIOID CONFIRMATION [254043] will automatically be reflexed after result verification. Send sample to RRL for confirmatory testing. |
| Patient is ≥ 14 years oldAND\*Patient is not a new mother | POSITIVE. Confirmation upon request only. |

 |

Updated 2/17/2023 LT