Overview

| Page | Topic |
| --- | --- |
| 3 | • Purpose |
| 3-4 | • Materials • Reagents • Supplies • Equipment |
| 4-5 | • Sample Information: [Specimen Type, Specimen Rejection Criteria, Specimen Volume, Specimen Stability] • Safety Precautions |
| 5 | • Quality Control |
| 6 | •Running QC: iChemVELOCITY analyzer |
| 7 | • Running QC: iQ200 analyzer |
| 8 | • Preparing Patient Specimens • Logging on to the System |
| 9 | • Operating the iRICELL System |
| 9-10 | • Locating Auto-Released results • Performing On-Screen Verification of Results |
| 12 | • Assaying Diluted Specimens |
| 14 | • Performing a Search |
| 15-16 | • Accessing Consumables Traceability • Handling an Expired Consumables Alarm |
| 16-17 | • Changing the lot number and expiration date of Chemistry QC • Utilizing the HELP menu |
| 17 | • Enabling the Detailed Audit Trail |
| 18 | • Interpretation/Results/Alert values |
| 19 | • Reference Intervals |
| 20-22 | • Method Performance Specifications • Result Reporting • References • Related Procedures • Appendices |

|  | Supplemental Material |
| --- | --- |
| 22 | Appendix A: Theory of Operation and Principle |
| 25 | Appendix B: Clinical Significance of Urine Chemistry Results |
| 27 | Appendix C: Limitations and Interferences |
| 29 | Appendix D: Expected Results |
| 30 | Approval Signatures |

| # Purpose |  | This procedure template provides instructions for the Iris Automated Urinalysis: iRICELL. |  |  |
| --- | --- | --- | --- | --- |
| Materials/ Equipment | Consumables and Part Number |  | Storage | Packaging and Use |
| iChemVELOCITY Consumables |  |  |  |  |
|  | iChemVELOCITY Strips [800-7212] |  | Room temperature. Stable unopened until expiration date on bottle. Open vial stability=5 days | Bottle of 100 strips; 1 each. Use daily. |
|  | iChem Wash Solution and Wash Filter [800-7704] |  | Room temperature. Stable unopened until expiration date on bottle. Open bottle stability=3 months | 4 bottles/case; 5.5ml/test. Use daily. |
|  | IRISpec™ CA/CB/CC [800-7702] |  | Refrigerate 2-8 °C. Unopened stability= date on bottle. Bring poured aliquot to room temperature before use. Open stability= 15 days. Record open date on bottle. | 9 bottles (3 bottles of each control); 2 ml of each per day. Use daily. |
|  | iChemVELOCITY CalChek Kit: Strips [800-7703] |  | Room temperature Stable unopened until expiration date on bottle. Open stability=CalChek strips are single use only. | 2 vials/kit; 1 vial /quarter; single use. Use quarterly. |
|  | iChemVELOCITY CalChek Kit: Reagents [800-7703] |  | Room temperature Stable unopened until expiration date on bottle. Open tube stability=8 hours | 2 sets/kit; 1 set/quarter; single use. Use quarterly. |

| Materials/ Equipment | Consumables and Part Number | Storage | Packaging and Use |
| --- | --- | --- | --- |
| iQ200 Consumables |  |  |  |
|  | IQ® Lamina [800-3102] | Room temperature. Unopened stability= date on bottle. Change filter when replacing first bottle from a new case of 4. | 4 bottles/cs; 7000 mL/bottle; 14mL/test. Use daily. |
|  | Iris System Cleanser [800-3203] | Room temperature Unopened stability= date on bottle. | 4 bottles/cs 475 mL/bottle; 3ml/test. Use daily. |
|  | Iris Diluent [800-3202] | Room temperature Unopened stability= date on bottle. | 4 bottles/cs 475 mL/bottle; 3 ml/test. Use daily. |
|  | iQ® Calibrator [800-3103] |  Refrigerate 2-8 °C.  Unopened stability= date on bottle. Open bottle stability=24 hours | 4 bottles/cs 125 mL/bottle; 2ml/test. Use monthly. |
|  | iQ® Control/Focus Set [800-3104] |  Refrigerate 2-8 °C. Unopened stability= date on bottle. Open bottle stability=30 days | 1 bottle ea Pos/Neg; 2 bottles of Focus; 125 mL/bottle; Lot-specific Barcodes: Positive control, Negative control; Focus reagent; 3ml of Control/test; 6 ml of Focus/test. Use daily. |
|  | Dilution Code labels [800-3211] | N/A | 1 set; 1 each /test as needed. Use as needed. |
|  | 16x100mm polystyrene tubes | N/A |  |

| Sample | Specimen Type: - A clean-catch urine specimen should be obtained. If microbiology studies are ordered the specimen must be collected in a sterile container. The specimen must be labeled according to laboratory policy on the side of the container. Specimens with labels on the lid will be discarded. Specimen Rejection Criteria: - Reject grossly bloody specimens for the iChemVELOCITY chemistry analysis; dilute grossly bloody specimens for iQ200 microscopy analysis. Specimen Volume: - The specimen volume placed on the iRICELL must be at least 4 mL. - If testing on the iChemVELOCITY urine chemistry analyzer alone, the minimum volume is 2 ml. - If testing on the iQ200® urine microscopy analyzer alone, the minimum volume is 3 ml. Specimen Stability: - Urines, kept at room temperature, are stable for two hours. - If the specimen is not processed within two hours after collection, cap the container tightly and store at 2 - 8° C. - Specimens must be at or brought to room temperature before analysis. - Specimens collected in BD vacutainer UA preservative tubes are stable for 72 hours at room temperature. |
| --- | --- |
| Safety | - All reagents and controls should be handled as though capable of transmitting infectious diseases. Wear appropriate personal protective equipment when running patient samples or performing scheduled maintenance. - Refer to: Policy and Procedures Safety Manual Infection Control and Procedures 11-030-01 |

| Quality Control | Assay QC once every 24 hours for both chemistry and microscopy modules. QC material for microscopy module: - iQ® Positive Control - iQ® Negative Control - iQ® Focus reagent. QC material for chemistry module: - IRISpec CA/CB/CC controls Dipsticks: - Sticks are checked for reactivity and accuracy once every 24 hours by running the CA/CB/CC control material. - Test control material additionally when a new shipment/new lot number of iChemVELOCITY test strips is received. Record results in parallel testing log in preventive maintenance book. Parallel testing between the old shipment or lot number and the new shipment or lot number ensures that the system is operating within acceptable criteria. Upper limits, lower limits and target values are encoded within the lot-specific barcodes for the Positive, Negative and Focus iQ control material. - Criteria for Acceptable Control Results: - Quality control material results must fall within the ranges provided by the manufacturer of the material. - Refer to the section below for corrective action(s) if the values fall outside of the published limits. Note: The system will “lock out” patient testing for microscopy when microscopy controls fail and patient testing for chemistry when chemistry controls fail. |
| --- | --- |

| Running QC: iChemVELOCITY system | Follow the activities in the table below to run QC on the iChemVELOCITY. |  |  |  |
| --- | --- | --- | --- | --- |
| Before you begin | - Make sure that only unexpired reagents are used - Expired reagents will cause an ALARM |  |  |  |
|  | Step | Action |  |  |
|  | 1 | - Remove an iChemVELOCITY control rack. - Pour 2 mL of the CA, CB and CC control into 3 separate 16x100mm glass (or polystyrene plastic) tubes. - Return tightly capped controls to the refrigerator immediately. - Allow aliquots to warm to room temperature. - Use aliquots within one hour of pouring. |  |  |
|  | 2 | - Place the CA control in position 8. [color coded on bottle and tube position as RED] - Place the CB control in position 9. [color coded on bottle and tube position as BLUE] - Place the CC control in position 10. [color coded on bottle and tube position as YELLOW] |  |  |
|  | 3 | Place the rack on the right side of the iChemVELOCITY sampler touching the edge nearest the operator. Note: The rack will automatically advance forward [blue MEASURE light will appear]. |  |  |
|  | 4 | Control results will automatically print on the Quality Review screen as CA, CB and CC. |  |  |
|  | 5 | Look at QC results. |  |  |
|  |  | If… | Then… |  |
|  |  | QC Failed | - Look at the result printout to see why the control failed. - Adjust positions, if necessary. - Re-pour and re-run all controls. - If results are still not acceptable – check strips loaded in instrument for discoloration. If discolored, discard and add new strips and repeat. - If results are still not acceptable, notify supervisor and/or Iris Support. |  |

| Running QC: iQ200 analyzer. | Follow the activities in the table below to Run QC on the iQ200 analyzer. |  |  |  |
| --- | --- | --- | --- | --- |
| Before you begin | - If control is refrigerated, bring control set to room temperature before use. - Make sure that only unexpired reagents are used - Expired reagents will cause an ALARM |  |  |  |
|  | Step | Action |  |  |
|  | 1 | - Place appropriate barcode labels on 16x100mm glass (or polystyrene plastic) tubes [Positive, Negative, Focus]. - Mix iQ Positive control and iQ Focus by holding the bottle upside down and giving five hard sharp shakes followed by five gentle inversions. |  |  |
|  | 2 | Pour the following into 16x100 mm glass (or polystyrene plastic) tubes and place in the iQ200 QC rack: - Position 1: 3ml of Iris Cleanser - Position 2: 3ml of Iris Diluent - Position 3: 3ml of Iris Diluent - Position 4: Leave Empty - Position 5: 6 ml of Focus reagent - Position 6: 3 ml of Positive Control - Position 7: 3 ml of Negative Control Control rack must be run immediately after pouring focus and controls. |  |  |
|  | 3 | - Place the rack on the iQ200 sampler. - Press the START button located at the upper left corner of the iQ200, if the instrument is in the standby mode (green light). - If in measure mode (blue light), the rack will be detected and processed automatically. |  |  |
|  | 4 | Review results under: - Quality Review screen - QC Statistics |  |  |
|  |  | If… | Then… |  |
|  |  | QC Failed | - Look at message code to see why the control failed. Note: If control failed due to an identification error or QC out of order, resolve this error. - Pour fresh aliquots and re-run control. - If results are still not acceptable, notify supervisor and/or Iris’ Clinical Support. |  |

| Preparing Patient Specimens: | Follow the activities in the table below to Prepare Patient Specimens |  |  |
| --- | --- | --- | --- |
|  | Step | Action |  |
|  | 1 | Bring Samples to room temperature |  |
|  | 2 | Label an empty 16 x100mm glass tube (or polystyrene plastic) with patient identifier or use previously labeled urine preservative tube for testing. - Apply the barcode to the tube so that the start of the barcode (not the label edge) is approximately ½ inches from the top of the tube. Note: This leaves room for the dilution label should it be required. |  |
|  | 3 | Mix sample thoroughly by inversion. |  |
|  | 4 | Pour 4ml of well-mixed specimen into labeled tube. |  |
|  | 5 | Put the sample tube in position number 1 on the sample rack Note: - The rack’s black barcode should be facing to the right. - The tube’s barcode should face the instrument. |  |
|  | 6 | Load up to 10 samples in each rack in consecutive positions. |  |

| Logging on to the iRICELL: | Follow the activities in the table below to Log on to the iRICELL: |  |  |
| --- | --- | --- | --- |
|  | Step | Action |  |
|  | 1 | - Access the Instrument Screen. - Click on the Logon button. |  |
|  | 2 | - Click on the down arrow. - Select your identifier (Do Not type in your identifier). - Press tab. |  |
|  | 3 | Enter your password. |  |
|  | 4 | Click on OK. Note: The current user’s name will now appear at the top of the Instrument screen. |  |

| Operating the iRICELL System | Follow the activities in the table below to Operate the iRICELL System |  |  |
| --- | --- | --- | --- |
|  | Step | Action |  |
|  | 1 | Ensure that sufficient supplies and consumables are loaded. |  |
|  | 2 | Load up to 10 samples in each rack in consecutive positions. |  |
|  | 3 | Place rack(s) on right hand side of the iChemVELOCITY. - Ensure that the notch of the rack base is placed onto the Sampler track ridge. Note: The instrument will automatically advance the rack. |  |
|  | 4 | THE REMAINDER OF THE PROCESSING IS PERFORMED AUTOMATICALLY ON THE SYSTEM: - The sample rack will be moved along the sample transport tray to the barcode reader. - After the barcode is read, the probe mixes the sample, aspirates an aliquot, analyzes the SG, color, clarity and dispenses the sample onto a test strip. - When the sample processing is complete, the sample rack will be automatically transferred, via the bridge, to the iQ200 Analyzer. - If the specimen does not need to be run on the iQ200 Analyzer, remove the rack from the instrument at this point. If the rack is to be run on the iQ200 Analyzer, allow the rack to transfer across the bridge. - After the rack is transferred via the bridge to the iQ200 Analyzer, the sample rack will be moved along the iQ200 Sampler to the barcode reader. - The iQ200 Analyzer barcode reader reads the specimen barcode. - If a microscopic examination is to be done (as determined by the user-defined criteria), the probe will mix the sample, aspirate an aliquot and perform the microscopic examination. If a microscopic examination is not to be performed, the tube will be passed. - After sample processing is complete, unload the sample racks from the left side of the iQ200 Analyzer. |  |
|  | 5 | Note: Completed, auto-released results will appear on the Found List. Verify any pending, flagged results on the Work List as follows in next section of this document. |  |

| Locating Autoreleased Results | Follow the activities in the table below to Locating Autoreleased Results Note: - The majority of results will have been auto-released. - Results can be printed to the printer, the LIS or the screen |  |  |
| --- | --- | --- | --- |
|  | Step | Action |  |
|  | 1 | Click on the Search button to access the Found List |  |
|  | 2 | Specimen results will appear as user defined. |  |
|  | 3 | No further action is necessary |  |

| Performing On-Screen Verification of Results | Follow the activities in the table below for Performing On-Screen Verification of Results Note: - If the auto-release feature has been enabled, the majority of results will have been auto-released. - Verify results only for those specimens not auto-released [i.e. specimens appearing on the Work List]. - Verify only Yellow-colored categories - Verify using FULL EDIT |  |  |
| --- | --- | --- | --- |
| Before You Begin | - Green and Red colored categories will be autoreleased (or require no review) - Yellow categories require on-screen review |  |  |
|  | Step | Action |  |
|  | 1 | Click on the Work List button located on the top right part of the instrument screen to bring up all unreleased samples. Note: Samples may be sorted by Specimen ID, Date-Time, Rack/Pos or Status by header at the top of the row. (Clicking a second time will reverse the order.) |  |
|  | 2 | Highlight the specimen to be verified as follows: - Click on Specimen Identifier: - Click on the Specimen button. - Verify consolidated Chemistry and Microscopy results. |  |
|  | 3 | Clear flags that are displayed before results are verified or deleted as follows: - To Review Flagged Specimen: - Click Review Flagged Specimen button. - Click Accept. - To Delete Specimen results (if the result must be discarded): - Click Delete Flagged Specimen button. - Click Accept. |  |

|  | 4 | Verify auto-classified particles as follows: - Click on “Edit” Note: You will be directed to the first yellow particle category |  |  |
| --- | --- | --- | --- | --- |
|  |  | If | Then |  |
|  |  | The classification of the particle is acceptable, | Continue the verification by clicking on the right arrow to move forward. Note: The Yellow category will appear. |  |
|  |  | The classification is not acceptable, | Reclassify the misidentified particle(s) as follows: - Determine whether or not reclassification will make a clinical difference. - Reclassify particles only when it will make a clinical difference. - Click on the particle type that the image(s) should be classified into (use right-hand button). - Click on the image(s) to be moved. - Press the forward arrow to proceed to the next category. Note: - If all images of a category are misclassified, click on the particle type and then click on the right arrow to move to the next category. |  |
|  | 5 | Press ACCEPT to release the results when verification is complete. |  |  |
|  | 6 | Return to the microscope for the following: - Oval fat bodies (to view using polarized light) - Fat (to view using polarized light) - Trichomonas (to observe motility) - Cellular Casts (only necessary when the operator cannot make a definite identification of the cell type using the iQ) |  |  |

| Assaying Diluted Specimens | Follow the activities in the table below to Assay Diluted Specimens |  |  |  |
| --- | --- | --- | --- | --- |
| Before you Begin | Dilutions must be prepared for the iQ for the following specimens: - Grossly bloody - Mucoid - Dense/Viscous - Short Samples Note: - Dilutions are run on the iQ200 urine microscopy analyzer only - Do Not perform dilutions for the iChemVELOCITY urine chemistry analyzer. - Identify specimens that require a dilution before placing specimen on the iRICELL. |  |  |  |
|  | Step | Action |  |  |
|  | 1 | Select dilution and corresponding dilution code (refer to Dilutions under Formed Particles or Fluid Type under Settings on the Instrument Screen). |  |  |
|  |  | If… | Then… |  |
|  |  | You have barcodes | - Obtain a Dilution Rack - Fix identical patient barcode onto two unlabeled tubes. - Pour 3 ml urine into the first tube. - Place the tube in the dilution rack and run on the Chemistry analyzer. Note: The rack will proceed to the iQ, but will not be aspirated. - Remove the rack from the Chemistry analyzer. - Label the matching second tube with the appropriate secondary barcode dilution label (fix label below the patient barcode). - Prepare the dilution, in this tube, using Iris Diluent. - Replace the undiluted sample that was used for the Chemistry analyzer with the diluted tube and place the specimen into a patient rack. - Put rack on the iQ analyzer - Press START to run the sample. - Results will consolidate. - If auto-release has been enabled, results will auto-release as the user has defined, unless flagged. - Verify results only for non-autoreleased samples (refer to previous sections of this document). |  |

| Assaying Diluted Specimens continued… |  | If continued… | Then continued… |  |
| --- | --- | --- | --- | --- |
|  |  | You do not have barcodes | - Obtain a dilution rack - Click on Manual Orders - Choose the patient rack and position number that will be used to run the assay. - Identify the specimen ID, select URN, Dilution code and Work order=run. - Put the sample into the correct position in Dilution rack number 23. - Pour 3 ml urine into the corresponding unlabelled tube. - Place rack on the right hand side of the Chemistry analyzer and run. - [Note: the Dilution rack will not be aspirated by the iQ] - The Chemistry Result will be displayed as ID_ERROR. - After Chemistry has completed, remove rack. - Perform appropriate dilution. - Place the diluted sample in the patient rack and position number that was programmed [refer to second bullet). - Place the rack on the iQ analyzer. - Press START to run the sample. - Consolidate results. - If auto-release has been enabled, results will auto-release, as the user has defined, unless flagged. - Verify results only for non-autoreleased samples (refer to previous sections of this document). |  |

| Performing a Search | Follow the activities in the table below to Perform a Search. |  |  |  |
| --- | --- | --- | --- | --- |
|  | Step | Action |  |  |
|  | 1 | - Access the Work List List - Select SEARCH |  |  |
|  | 2 | Clear all previous entries by Clicking “Clear”. |  |  |
|  |  | To Search by | Then |  |
|  |  | Specimen ID | - Enter specimen identifier - Press OK |  |
|  |  | Sequence number | - Enter desired Sequence number - Press OK |  |
|  |  | Operator | - Enter Operator Log-in ID - Press OK |  |
|  |  | Date and Time | - Enter the specific date and time - Press OK |  |
|  |  | 24 hours | - Select 24 hours - Press OK |  |
|  |  | Today | - Select Today - Press OK |  |
|  |  | Lot | - Select Lot number - Select specific lot from lots displayed - Select OK [Note: This function enables the user to search for Results Tied to a Specific Lot number] |  |
|  |  | Last Name | - Enter the Last Name or range of last names. - Press OK |  |
|  |  | First Name | - Enter First Name or range of first names. - Press OK |  |
|  |  | Age | - Enter in the specific age expressed in decimals (For example: 10 years and 5 months is entered as 10 and 5/12=10.42) - Press OK |  |
| Continued on next page |  |  |  |  |

| Performing a Search | Step | Action continued… |  |  |
| --- | --- | --- | --- | --- |
|  |  | Location | - Enter Location - Press OK Note: Location must match the LIS entry exactly. |  |
|  |  | Urine Additional Criteria | - Enter Urine Culture Candidates or No Urine Culture Indicators Observed. - Press OK |  |
|  |  | Specimens awaiting transmission | - Check “Show Specimens Awaiting Transmission Only” and only these specimens will appear - Press OK. |  |
|  |  | Released specimens | - Check “Show Released Specimens Only” and only released specimens will appear. - Press OK Note: Imported images appear as Released specimens. |  |
|  |  | Incomplete Specimens | - Check “Show Incomplete Specimens Only” and only incomplete or specimens in progress will appear. - Press OK |  |

| Accessing Consumables Traceability | Follow the activities in the table below to Access Consumables Traceability |  |
| --- | --- | --- |
|  | Step | Action |
|  | 1 | Log in as Manager |
|  | 2 | Go OFFLINE |
|  | 3 | Access Consumables |
|  | 4 | Select Traceability |
|  | 5 | View all reagents that have been automatically entered by the reagent’s barcode |
|  | 6 | Manually enter reagents that do not have barcodes as follows: - Enter Reagent expiration date, start date, etc. - Select ADD - Select Update |
|  | 7 | Delete manually entered reagents as follows: - Enter a Deletion Comment - Select Update Note: The reagent will not disappear completely but will be removed from the queue. |

| Handling an Expired Consumables Alarm | Follow the activities in the table below to Handle an Expired Consumables Alarm |  |
| --- | --- | --- |
| Before you begin | If an expired reagent or strip is run on the instrument, a red alarm will appear. |  |
|  | Step | Action |
|  | 1 | If the alarm indicates that an expired Chemistry control or strip was run: - Check under Chemistry QC to make certain that the correct information has been added. - Correct information if incorrect (refer to Changing the lot number and expiration date of Chemistry QC) - Re-run control or strip |
|  | 2 | - If the alarm was due to another consumable: - Obtain an unexpired consumable - Re-run |

| Changing the lot number and expiration date of Chemistry QC | Follow the activities in the table below to Change the lot number and expiration date of Chemistry QC |  |
| --- | --- | --- |
|  | Step | Action |
|  | 1 | Select Consumables |
|  | 2 | Select Chemistry QC |
|  | 3 | Enter Strip lot and Expiration date listed on the urine chemistry strips |
|  | 5 | Select Next |
|  | 6 | Enter Chemistry control CA lot and expiration date listed on the reagent box |
|  | 7 | Select Next |
|  | 8 | Enter Chemistry control CB lot and expiration date listed on the reagent box |
|  | 9 | Select Next |
|  | 10 | Enter Chemistry control CC lot and expiration date listed on the reagent box |
|  | 11 | Select OK |

| Utilizing the HELP menu | Follow the activities in the table below to Utilize the HELP menu |  |
| --- | --- | --- |
|  | Step | Action |
|  | 1 | Access Instrument Screen |
|  | 2 | Select “?” icon |
|  | 3 | Operator’s Manual will appear as a PDF |
|  | 5 | Click on PDF to open the manual |
|  | 6 | Find topic desired and click on Table of Contents Note: The operator will be directed to the appropriate section |

| Enabling the Detailed Audit Trail | Follow the activities in the table below to Enable the Detailed Audit Trail Note: - The detailed Audit Trail will capture any change made to the system and record the name of the Operator who made the changes - Enabling this function required to Edit Chemistry results |  |
| --- | --- | --- |
| Before You Begin | Do Not Enable this Option without consulting with an Iris Product Specialist or Clinical and Customer Care Specialist before altering a setting. |  |
|  | Step | Action |
|  | 1 | Log on as a Manager |
|  | 2 | GO OFFLINE |
|  | 3 | Select SETTINGS |
|  | 4 | Select Specimen |
|  | 5 | Check Enable Detailed Audit Trail |

| Interpretation/ Results/Alert values | INTERPRETATION OF RESULTS: - Glucose, protein, and ketones are reported with semi-quantitative grading and as mg/dL. - Bilirubin is reported with semi-quantitative grading and mg/dL - Leukocyte esterase is reported as positive or negative. - Urobilinogen is reported with semi-quantitative grading and as mg/dL - pH is reported in quantitative units. - Blood is reported semi-quantitatively and as mg/dL. - Nitrite is reported as positive or negative. - Specific Gravity is reported by refractive index quantitatively with a value to 3 decimal places, ranging from 1.000 to >1.060. in 0.001 increments. - RBCs and WBCs are reported as cells per HPF. - Renal, transitional cells are numerated per HPF. - Squamous epithelial cells are reported per HPF. - Bacteria are usually reported qualitatively. - Crystals are reported qualitatively per HPF. - All casts are reported by type and enumerated per LPF. - Yeast is reported as none or present. - The presence of trichomonas must be verified by microscopy and the presence of motility. |
| --- | --- |
| Reference Intervals | ## Chemistry Results Specific Gravity 1.005 – 1.030 pH 5.0 – 8.0 Leukocyte Esterase Negative Nitrite Negative Protein, Qualitative <30 mg/dL Glucose Negative Ketones Negative Urobilinogen Negative Bilirubin Negative Blood Negative ## Microscopy Results WBC 0-5 /HPF RBC 0-3 /HPF Bacteria None Squamous Epi’s 0 /HPF Transitional Epi’s None Renal Epi None # Casts None Hyaline Casts None Broad Casts None Granular Casts None Cellular Casts None WBC Cast None RBC Cast None Waxy Cast None Fatty Cast None Epi Cell Cast None # Crystals None Yeast None WBC Clumps None Oval Fat Body None Trichomonas None Mucus None |
| Method performance specifications | pH is measured from 5.0 to 9.0 in 0.5 increments. Specific Gravity is measured via refractive index from 1.000 to >1.060 in 0.001 increments. Instrument linearity for microscopic particles is from 0-1000/uL, 0-182/HPF, or 0–2857/LPF. *Refer to iChemVELOCITY Urine Chemistry Strips for dipstick analytical sensitivity and report ranges. *Refer to Limitations and Interferences in the Appendix C of this procedure. |
| Result reporting | - Microscopic exams are performed on all urines with a positive leukocyte esterase, blood, nitrite and protein (greater than trace result). - Iris results for the urine elements that meet the criteria will auto-release to the LIS. Results that are above the minimum threshold or below the maximum threshold will not be considered for the auto-release and the results will be reviewed and released by the CLS. Refer to the autoverification criteria rules for urine elements below. - Perform verification of results as listed above for those not auto-released. - Glucose results that are greater than or equal to 500 mg/dL on patients under the age of 18 must be verified by manual dipstick testing. Users will be notified via a “popup” in the IQ 200 resulting screen if this condition exists. If the glucose result from manual testing is greater than or equal to 1000 mg/dL, the result is considered critical and the provider must be notified and this documented in the LIS. Note: Results are released to the LIS (or middleware) when the operator clicks on the “Accept” button. Analyte; Result criteria; Rule. Specific Gravity pH Glucose Ketones Urobilinogen Bilirubin Leukocyte Esterase Blood Nitrite Protein; Any value Any value Any value Any value Any value Any value Negative Negative Negative Negative Equal or less than trace; Report, “Micro exam not done” Or “Micro exam not indicated”. And auto-release results. Specific Gravity pH Glucose Ketones Urobilinogen Bilirubin Leukocyte Esterase Blood Nitrite Protein; Any value Any value Any value Any value Any value Any value >0 >0 >0 >Trace; Perform micro exam If any result meets criteria. Urine Microscopy Analyte; Result criteria; Rule. RBC/hpf WBC/hpf SQEP/hpf Results not meeting the criteria for auto-release will be reviewed by testing personnel.; If <4 or > 25 If <6 or > 25 If <1 or >10; Auto-release Auto-release Auto-release. |
| References | - iChemVELOCITY Operators Manual - Iris iQ200 Operators Manual - iChemVELOCITY urine chemistry strip insert - Fundamentals of Urine and Body Fluid Analysis, Nancy A. Brunzel, 2nd edition, 2004. - Urinalysis and Body Fluids, Susan King Strasinger, 5th edition, 2008. - GP16-A2: CLSI Urinalysis and Collection, Transportation, and Preservation of Urine Specimens; Approved Guideline-Second Edition |
| Related procedures | - Iris Automated Urinalysis: iRICELL™ System Maintenance Procedure |
| Appendices | - Theory of Operation. - Clinical Significance. - Limitations and Interferences. - Expected Values. |

APPENDIX A

PRINCIPLE AND THEORY OF OPERATION:

INTENDED USE:

The iRICELL™ is an in-vitro diagnostic system. Distinguished by throughput, the workcell includes the iRICELL 2000 (iQ 200 ELITE™ plus the iChemVELOCITY) and the iRICELL 3000 (iQ 200 SPRINT™ plus the iChemVELOCITY). The workcell is composed of the Automated Urine Chemistry Analyzer module [The iChemVELOCITY], the iQ Series Automated Urine Microscopy Module [iQ Series], results/analysis processor, computer monitor, mouse and keyboard. The iRICELL provides a fully integrated, automated chemical and microscopic analysis of urine. This procedure provides instructions for performing a complete urinalysis using the iRICELL.

THEORY OF OPERATION:

The iChemVELOCITY performs the chemistry panel, determines the specific gravity, color and clarity of a urine specimen. The chemistry panel is performed using a test strip, which detects the presence of 9 elements – glucose, protein, bilirubin, urobilinogen, pH, blood, ketones, nitrite, and leukocytes by wavelength reflectance. Specific gravity is determined by measuring the refractive index. Color is measured by transmitted light and clarity is measured by scattered light. The iQ Series performs the microscopic portion of the urinalysis and provides a quantitative or qualitative count of formed elements such as cells, casts, crystals, and organisms. The iQ Series photographs particles as they are passed in front of a digital camera. The images are classified, counted and stored for verification by the user. The workcell consists of a computer that is interfaced with an approved chemistry analyzer and the iQ Series modules. At the workcell, results of the chemistry profile and the microscopic analysis are collated, compared to user-defined criteria for auto-release, and stored for auto-release or verified. The majority of specimens can be autoreleased if user-defined criteria are entered. The user can verify results including the images of the formed elements. As needed, the user may sub-classify or verify results. After verification the results may be sent to the host computer or printed.

PRINCIPLE: iChemVELOCITY URINE CHEMISTRY SYSTEM

The iChemVELOCITY is a urine chemistry analyzer that measures the chemical constituents of the urine using iChemVELOCITY strips, which are read by a dual wavelength reflectance system. The iChemVELOCITY strips consist of a plastic strip containing nine (9) pads impregnated with chemicals specific for the determination of a particular constituent. The nine analytes measured are: glucose, protein, bilirubin, urobilinogen, pH, blood, ketones, nitrite, leukocytes esterase, ascorbic acid and a color compensation pad. A color compensation pad is included on the strip to compensate for the natural color of urine and its effect on the color of the reaction pads. Test strips are placed onto a strip conveyor system by a mechanical extractor. The sample probe mixes the sample, aspirates an aliquot of urine and dispenses it onto each reagent pad. At defined wavelengths, the iChemVELOCITY analyzes the color changes and the intensity of reflected light from the reaction pads. These measurements are used to calculate clinically meaningful results.

PRINCIPLE: iQ200 MICROSCOPY SYSTEM

The microscopic portion of a routine urinalysis is performed on the iQ200 Analyzer. The iQ200 Analyzer auto-identifies and processes specimens by mixing, sampling and analyzing the data obtained from the sample. Approximately 1mL of the mixed specimen is aspirated and is sandwiched between enveloping layers of a suspending fluid. This fluid or “lamina” is positioned exactly within the depth of focus and field of view of the objective lens of a microscope that is coupled to a video camera. The iQ® Lamina is used to position the formed elements in the best orientation that presents the particles with their largest profile facing the direction of view. The camera captures five hundred pictures per sample. The flash of a strobe lamp illuminates each field. The pictures are digitized and sent to the instrument processor. Individual particle images are classified into one of 12 categories using size, shape, contrast and texture. The auto-classified categories are RBCs, WBCs, WBC clumps (WBCC), hyaline casts, unclassified casts (UNCC), squamous epithelial cells, non-squamous epithelial cells (NSE), bacteria, budding yeast, unclassified crystals (UNCX), mucus, and sperm. Any images that do not classify into any one of these 12 categories are placed in the UNCL category. The particle concentration is calculated using the number of images, normalization factor and the volume scanned. User defined criteria for the auto-release of results is checked and results are sent either directly to the host computer and/or printer or to the system monitor for verification. Only non-autoreleased specimens will appear on the system screen for verification. During the verification process, individual images may be displayed. The operator may manually reclassify digital images when necessary and in accordance with the laboratory’s policy. Unclassified crystals (UNCX), unclassified casts (UNCC), and non-squamous epithelial cells (NSE) may be further sub-classified during the verification process. XXX State here if there are any criteria for your facility that must be met to allow results to be reported. (Example: If XXX are not sub-classified, the accession will fail and cannot be verified in the LIS.) Once the verification process has been completed and “ACCEPT” has been chosen, the results will be sent to the LIS, screen or printer.

iChemVELOCITY DIPSTICK CHEMICAL REACTIONS:

| Bilirubin | This test is based on the coupling of bilirubin and diazonium salt in an acidic medium. A pinkish tan color proportional to bilirubin concentration is generated. |
| --- | --- |
| Urobilinogen | This test is based on the coupling reaction of urobilinogen with a stable diazonium slat in buffer. A pink to red color proportional to the urobilinogen concentration is generated. |
| Ketone | This test is based on Legal’s method in which the test pad contains sodium nitroprusside and glycine in an alkaline medium. A violet color proportional to methylketone is generated. |
| Ascorbic Acid | This test is based on Tillman’s reaction in which the presence of ascorbic acid leads to the decolorization of the pad from gray-blue to orange. |
| Glucose | This two-step enzymatic reaction uses glucose oxidase, peroxidase and a chromogen. Glucose oxidase catalyzes the formation of gluconic acid and hydrogen peroxide via the oxidation of glucose. Peroxidase then catalyzes the reaction of hydrogen peroxide with a chromogen via the oxidation of chromogen to colors ranging between green and gray-blue. |
| Protein | This test is based on the “protein error of pH indicators” on the green color developed from the presence of protein. This dye-binding is particularly strong with albumin. |
| Blood | This pseudo-enzymatic test contains organic peroxide and a chromogen. The peroxidase effect of hemoglobin and myoglobin causes a color change to green. |
| pH | This test contains a mixed indicator which assures a marked change in color between pH5 and pH9. Colors range from orange through yellow and green to cyan. |
| Nitrite | This test is based on modified Griess reaction in which nitrite in the urine reacts with amide to form a diazonium compound. The subsequent coupling reaction yields a pink color in the presence of nitrite. Some Gram positive and non nitrite-forming bacteria are not detected in this test. |
| Leukocytes | This enzymatic test pad contains an indoxyl ester and a diazonium salt. Granulocyte esterases react with indoxyl ester and diazonium salt to generate a violet color. |

APPENDIX B

CLINICAL SIGNIFICANCE OF URINE CHEMISTRY RESULTS:

| Glucose: | The presence of glucose in the urine called glucosuria is caused by hyperglycemia or renal condition. Diabetes mellitus is the most common disease resulting in hyperglycermia. Renal conditions causing dysfunction of tubular reabsorption of glucose occur in many conditions including pregnancy. |
| --- | --- |
| Protein: | The presence of protein in urine is often the first indicator of renal disease, but its appearance in the urine doesn’t always signify renal disease. Although proteinuria may indicate nephritic syndrome, multiple myeloma, glomerulonephritis, and pre-eclampsia, a transient mild proteinuria can be present after exposure to cold, strenuous exercise, high fever, dehydration, or an acute phase of a severe illness. The strip is primarily sensitive to albumin. |
| Bilirubin: | The appearance of urinary bilirubin can be a sign of liver disease or extra-or intra-hepatic biliary obstruction. |
| Urobilinogen: | The normal urine has a small amount of urobilinogen (less than or equal to 2.0 mg/dL). The strip is unable to detect a decreased amount, which may appear in infants, patients on antibiotic therapy, or patients with obstructive disease. Increased amounts appear in hemolytic anemias and liver dysfunction. |
| pH: | Along with the lungs, the kidneys are the major regulator of acid-base balance. Freshly voided urine has a pH of 5.0 – 6.0. The pH of urine can be controlled by dietary regulation and medication. |
| Blood: | A positive reaction for blood may indicate red cells, hemoglobin, or myoglobin present in the urine. Hematuria can be seen due to bleeding as a result of trauma or irritation (renal calculi, glomerulonephritis, tumors, toxic or chemical exposure). Hemoglobinuria occurs when there is lysis of red cells in the urinary tract, intravascular hemolysis or transfusion reactions. Very dilute or extremely alkaline urine can also lyse the cells. Myoglobinuria indicates muscular destruction that may appear in hypothermia, convulsions, and extensive exertions. |
| Ketones: | Ketonuria appears when there is an increased use of fat instead of carbohydrate as a source of metabolism. Conditions of ketonuria include diabetes mellitus, vomiting, inadequate intake of carbohydrates due to starvation, weight reduction, or pregnancy. |

APPENDIX B continued…

CLINICAL SIGNIFICANCE OF URINE CHEMISTRY RESULTS:

| Nitrite: | Bacteria, specifically gram negative organisms, are detected by this nitrite reducing reaction. In order for the reaction to take place there must be adequate dietary nitrates, and the urine must be in the bladder at least four hours for the bacteria to react with nitrate for a positive reaction. Unusually colored urine due to medication or dyes can interfere with this reaction. |
| --- | --- |
| Leukocytes: | The presence of white blood cells in the urine is an indicator of inflammation. Lysed and intact WBCs are detected because both may have produced esterase. |
| Specific Gravity: | Specific gravity is a measure of the dissolved substances present in the urine. Specific gravity is one measure of the concentrating and diluting ability of the kidneys and hydration status of the patient. The specific gravity is obtained by measuring the refraction angles of light passing through a triangle prism. An LED emits a beam of light through a slit and a lens. The refractive index changes according to the specific gravity of the sample, the higher the specific gravity the greater is the angle of measurement. The change in the angle of the light is reported as the specific gravity. The result is automatically corrected for elevated protein or glucose concentrations as measured by the test strip. |
| Color: | Color variation can indicate the presence of a disease process, metabolic abnormality, or an ingested food or drug, or the variation simply due to excessive physical activity or stress. The color of the specimen is measured by transmitted light. The colors obtained are colorless, yellow, orange, brown, red, violet, blue, green and other, including light and dark of each. |
| Clarity: | Substances that cause urine turbidity may be pathologic or non-pathologic. The clarity or turbidity of a urine specimen is measured by passing a beam of light through the sample and measuring how the light is scattered. The amount of scattered light increases as the specimen becomes more turbid. The amount of clarity is reported as clear, turbid or extremely turbid. |

APPENDIX C

LIMITATIONS AND INTERFERENCES:

| ANALYTE | CAUSES OF FALSE NEGATIVE RESULTS | CAUSES OF FALSE POSITIVE RESULTS |
| --- | --- | --- |
| Bilirubin | Elevated concentrations of nitrite may inhibit the reaction. Bilirubin is light sensitive and prolonged exposure of urine specimens to light may result in diminished or false negative values | Some urine specimens may contain impurities such as food dyes and therapeutic pigments to produce a yellowish or reddish discoloration of the test pad that may lead to the interference. Elevated Urobilinogen concentrations may slightly enhance the response to this test pad. |
| Urobilinogen | This test is inhibited by elevated concentrations of formaldehyde and nitrite ≥10 mg/dl. Prolonged exposure to light may lead to diminished or false negative values. | Food dyes and medications that have an intrinsic red color in acidic medium such as red beets, azo dyes, phenazopyridine and p-aminobenzoic acid may produce false positive results. |
| Ketones | Elevated concentrations of phenylpyruvic acid may interfere with the test pad and produce a variety of colors. Phthaleins and anthraquinone derivates exhibit a red color in alkaline medium and this may mask the response. Large amounts of levodopa and medications containing sulfhydrl groups may produce atypical color reactions. | N/A |
| Ascorbic Acid | No interferences reported | No interference reported. |
| Glucose | Ascorbic acid concentrations of up to 50 mg/dL did not interfere with glucose assay test results (no false negative results). Acetoacetic Acid concentrations of up to 200 mg/dL did not interfere with glucose assay test results (no false negative results). High specific gravity, acidic pH values and gentisic acid may inhibit color formation. | Cleaning agents such as hypochlorite and peroxide may lead to false positive results. |
| Protein | Food dyes such as red beets and therapeutic pigments such as methylene blue and pyridium may mask the coloration of the test pad. Interference may occur with high specific gravity. Interference may also occur with disinfectants, wetting agents and blood substitutes (quaternary ammonium compounds, polyvinylpyrolidone, chlorohexidine). |  |

APPENDIX C

LIMITATIONS AND INTERFERENCES continued…

| ANALYTE | CAUSES OF FALSE NEGATIVE RESULTS | CAUSES OF FALSE POSITIVE RESULTS |
| --- | --- | --- |
| Blood | Reducing agents such as ascorbic acid, uric acid, glutathione and gentisic acid may cause false negative results. Samples with a pH of 5 may interfere with this test. High concentrations of nitrite can delay the reaction. | Preservatives (formalin) and cleaning agents such as hypochlorite may result in false positives. |
| pH | No interferences reported. | No interferences reported. |
| Nitrite | A negative response in the presence of bacteriuria may be caused by the following: non-nitrite producing microorganisms, low nitrate diet, antibiotic therapy, strong diuresis, or insufficient urinary retention time in the bladder. | Food dyes and therapeutic pigments such as red beets and pyridium may cause false positive responses. |
| Leukocytes | High concentrations of protein, glucose, cephalexin and gentamicin may diminish the color response. The test can be negative in the presence of visible leukocytes if they have not lysed and/or are not granulocytes. | False positive results may occur in the presence of preservatives such as formaldehyde and formalin. Test results may be positive in the absence of observable cells if the granulocytes have lysed. |
| Specific Gravity | N/A-measured by Refractometer | N/A-measured by Refractometer |
| Color | N/A –measured by scattered light | N/A –measured by scattered light |

APPENDIX D

EXPECTED RESULTS:

| Bilirubin | In normal urine, no detectable level of bilirubin should be obtained. Positive results require further investigation. Performance Characteristics: The test is specifically developed for bilirubin. Biliverdin does not react with this test pad. |
| --- | --- |
| Urobilinogen | In normal urine, urobilinogen is usually present at concentration up to 1 mg/dl. A result of 2 mg/dl represents the transition from normal to abnormal state and the specimen should be further investigated for possible liver disease and hemolytic disorders. Performance characteristics: The diazonium-based test is more specific for urobilinogen than Ehrlich’s reagent based- test. Test strips cannot determine the absence of urobilinogen, which may be significant in biliary obstruction |
| Ketone | Detectable amounts of ketone do appear in the urine of normal specimens. Positive ketone values may result from the following conditions:starvation, dietary imbalance, diabetes mellitus, eclampsia, insulin dosage monitoring, vomiting and other metabolic disorders. Performance Characteristics: The test does not measure B-hydroxybutric acid and is only slightly sensitive to acetone. |
| Ascorbic Acid | Ascorbic acid is found in various food supplies and dietary supplements. Concentrations of ascorbic acid greater than 20 mg/dl can be expected to cause strong interference with glucose, blood, and nitrite. Performance Characteristics: The oxidized form, dehydroascorbic acid, does not react with this test pad. |
| Glucose | A small amount of glucose (up to 20 mg/dl) may be present in normal urine. The detectable sensitivity of this test has been adjusted to exclude the minute amount of glucose. Therefore, any positive response should be further evaluated. Performance Characteristics: The test pad doe not react with other reducing sugars such as fructose and galactose. |
| Blood | Normal urine contains no detectable hemoglobin or intact red blood cells. Any positive results should be further evaluated. Performance Characteristics: 0.015 mg/dl blood |
| pH | The normal pH of urine can vary between pH 4.5and pH 8.0. Performance Characteristics: pH values are determined to within 0.5 unit over the range from 5.0 to 9.0. |
| Nitrite | Normal urine contains no detectable nitrite. However, a negative result does not rule out a urinary tract infection. Performance Characteristics: This test is specific for nitrite. Results may depend on the ability of bacteria to reduce nitrate to nitrite, the number of bacteria and the retention time of the urine in the bladder. |
| Leukocytes | A normal urine specimen should not produce a positive result. The test can be negative in the presence of observable leukocytes if they are not lysed/and or are not granulocytes (example:lymphocytes). The test may be positive in the absence of observable cells if the granulocytes have lysed. Performance Characteristics: Leukocyte test detects the presence of esterase in the granulocytic white blood cells. The test result is most frequently accompanied by the presence of bacteria that may or may not produce a nitrite positive reaction. |

APPROVAL SIGNATURES

Effective Date

Effective date for this procedure:

Author

Compiled by Iris Diagnostics

Revised by:  Cindy Schwartz

Procedure Approved:

Date:    _____________________________________________________

Approved by:	 _____________________________________________________

Schedule for Review

Date Reviewed/Revised:_______ Approved:______________________Comments:

Date Reviewed/Revised:_______ Approved:______________________Comments:

Date Reviewed/Revised:_______ Approved:______________________Comments:

Date Reviewed/Revised:_______ Approved:______________________Comments:

Date Reviewed/Revised:_______ Approved:______________________Comments:

Date Reviewed/Revised:_______ Approved:______________________Comments:

Date Reviewed/Revised:_______ Approved:______________________Comments:

Date Reviewed/Revised:_______ Approved:______________________Comments:

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| Major | New Procedure to replace Urinalysis P&P 03-040 and 03-045 | Cindy Schwartz 07/29/14 |  |  |  |