#### POST VASECTOMY SEMEN ANALYSIS

| ##### PRINCIPLE | Postoperative semen analysis should be done to evaluate the vasectomy operation. The examination should note presence of sperm and motility. |
| --- | --- |

| ##### SPECIMEN COLLECTION | Specimen must be submitted as soon as possible as testing labs only. |
| --- | --- |

| ##### PROCEDURE |  |  |
| --- | --- | --- |
|  | STEP | ACTION |
|  | 1. | Record the time the specimen was collected and received in laboratory. |
|  | 2. | Motility must be checked within 30 min. of collection if possible. |
|  | 3. | Mix sample well. |
|  | 4. | Put one drop of semen sample on a clean slide using disposable pipette. |
|  | 5. | Cover with cover slip. |
|  | 6. | Observe under high power field (40x), scan at least 20 fields. |
|  | 7. | Note number of spermatozoa present and motility. |
|  | 8. | If no sperm are present, transfer semen sample to a conical centrifuge tube and centrifuge the sample at 2000 RPM for 3 minutes. |
|  | 9. | Decant the supernatant and re-examine the sample for spermatozoa and motility, examine 10 fields. |
|  | 10. | Report in Cerner using Accession Result Entry according to the chart below. |

| ##### RESULTS REPORTING | Reporting will be made in the increments according to the chart below. - Count is determined using the concentrated sample. - Motility is reported as present if seen in either the un-spun or the centrifuged specimen. |
| --- | --- |

| Observed sperm/HPF | Enter in Cerner |
| --- | --- |
| 0 | 0 |
| 1-3 | 0-3 |
| 4-20 | 4-20 |
| >20 | >20 |
| Motility | Motile |
| No Motility | Non-Motile |

| ##### REFERENCE RANGE | <4 PER HPF |
| --- | --- |

| ##### REFERENCE | Brownlee James H., Tibbels, Kelly C.: 'Vasectomy The Journal of amily Practice, Vol. 16, No. 2: 379384. 1983.Henry, John: Clinical Diagnosis and Management by Laboratory Methods: Vol 1, l5th Ed Examination of the Seminal Fluid, 712-719. Philadelphia, Saunders Company, 1979 |
| --- | --- |
|  | WHO Laboratory Manual for the examination of human semen and sperm-cervical mucus interaction, 4th ed.. Cambridge, UK: Cambridge University Press, 2000. |
|  | Baker, DJ.. Performing a Quality Semen Analysis in the Clinical Laboratory, MLO. December 2000. |

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| Major | Added concentration step | Cindy Schwartz 5/17/11 |  |  |  |
| Minor | Changed resulting method to Cerner appropriate choices | Cindy Schwartz 9/18/14 |  |  |  |

Imp. =Implemented