- COLD AGGLUTININ TREATMENT

| - PRINCIPLE | The presence cold agglutinin can adversely affect the results of the CBC, i.e. inaccurate results. The parameters affected are RBC, HCT, MCV, MCH and MCHC. This is due to the agglutination of the RBC's in the sample as it cools to room temperature. |
| --- | --- |

| - MATERIALS & REAGENTS | 37 ° C Drybath and timer Micro-hematocrit centrifuge Heparinized 75mm Hematocrit tubes Critoseal Cellpack Hematology Analyzer |
| --- | --- |

| - SAFETY | All specimens should be handled as though capable of transmitting infectious diseases. Wear appropriate personal protective equipment when running patient samples or performing scheduled maintenance. Refer to: Policy and Procedures Safety Manual Infection Control and Procedures 11-030-01. |
| --- | --- |

PROCEDURE

| Step |  |
| --- | --- |
| 1. | Incubate the CBC sample by putting it in a 37 ° C Drybath for at least 30 minutes. |
| 2. | Mix well and rerun the specimen immediately after pre-warming. |
| 3. | If MCHC is corrected, then the other indices have been corrected for cold agglutinin. Proceed to result the corrected values. |
| 4. | In cases of very strong cold agglutinins where the MCHC does not correct after warming the sample for more than 1 hour, perform a SPUN HCT. |
| 5. | If a manual diff is needed, prewarm several slides in the drybath for 15-30 minutes before making the smear. |
| 6. | Report only the WBC, HGB, SPUN HCT and PLT count. Add the comment: “Unable to report RBC and RBC Indices due to strong cold agglutinins. Hematocrit is a SPUN HCT.” |

| REFERENCE | Coulter Product Guide, July (1987), pp. 8-41. |
| --- | --- |

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Name of responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| Major | Updated the format, revised the principle, added the safety and materials/reagents section, revised the procedure. | Julius Salomon 5/9/2014 |  |  |  |

Imp. =Implemented