#### BODY FLUID CELL COUNT

| POLICY | Body fluid counts must be completed as soon as possible or within 1 hour of being received in the laboratory. |
| --- | --- |

| METHOD | Manual cell count |
| --- | --- |

| ##### SAFETY | All specimens, reagents and controls should be handled as though capable of transmitting infectious diseases. Wear appropriate personal protective equipment when running patient samples or performing scheduled maintenance. Refer to: Policy and Procedures Safety Manual Infection Control and Procedures 11-030-01 |  |
| --- | --- | --- |

| ##### MATERIALS & REAGENTS | - Microscope Cytology Funnel and Caps - Sterile Transfer Pipet Cytospin Centrifuge - Disposable C-Chip hemocytometer Cytology Clips |
| --- | --- |
| ##### SPECIMEN TYPE | The following types of specimen are considered to be body fluids: Pleural, Peritoneal, Gastric, Pericardial, and Synovial fluid. Note: See procedure 04-120 for Spinal Fluid Cell Count. |

| ##### SPECIMEN COLLECTION | Fluid specimens can be collected in sterile tubes containing anticoagulants for cell counts (EDTA is anticoagulant of choice), or large sterile containers. The large containers must be well mixed and a small portion (5 ml) aseptically removed and placed in a 5 ml EDTA tube for a cell count. Always use sterile technique when handling body fluid specimens. |
| --- | --- |

Note: Observe specimen for fibrin clots and pellicle formation. If small fibrin clots are detected, perform test as usual and include a comment in the report stating that results may not be accurate due to fibrin clot formation. If specimen is completely clotted, do not perform the cell count. Notify the nursing staff or provider that the specimen is clotted.

Continued on next page

PROCEDURE

GROSS EXAMINATION

| Step | Action |
| --- | --- |
| 1. | Note the physical appearance and check for fibrin clots. Appearance can be reported as clear, bloody, clotted, cloudy, hazy and slightly hazy. |

| ##### COMPLETING CELL COUNT | Follow the steps below to complete the cell count for RBC's and WBC's on body fluids. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Load both sides of hemocytometer chambers with an undiluted well mixed body fluid specimen. Note: On high counts a dilution may be necessary. Disposable C-Chip hemocytometer is for single use only, should be clean and free of scratches. |
| 2. | Using 40x magnifications, count the RBC's and WBC's in the appropriate number of large squares in both chambers. |
| 3. | The counts from each chamber must agree within 10% or the count must be repeated. Use the average of the two sides for the calculation. |
| 4. | Use the following formula to calculate number of cells/mm3: Number of cells counted x depth (10) x dilution = cells/mm3 Number of large squares counted NOTE: Cell counts and calculation of cell counts can be difficult. It is imperative that counts be calculated correctly so when in doubt consult with your co-worker or supervisor. |

| ##### PREPARING SLIDE | If the specimen contains White Blood cells, make a slide of the specimen by following the directions in Preparing Slides for the Cytospin procedure (Hematology P&P 04-190). |
| --- | --- |

Continued on next page

| ##### PERFORMING DIFFERENTIAL | Follow the steps below to complete the differential count on body fluids. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Count 100 WBC's, differentiating between neutrophil, lymphocyte, monocyte, eosinophil and basophil. |
| 2. | If blasts, atypical cells or unidentifiable cells are present, refer smear to Pathologists for review. Refer to Hematology Procedure 04-221 for procedural specifics. |
| 3. | The slides will be saved for a month. |

| ##### RESULTS REPORTING | Follow the steps below to complete the reporting of body fluid counts. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Results to be entered are: - Appearance: clear, bloody, clotted, cloudy, hazy, slightly hazy - Total RBC - Total WBC - WBC Differential (5 part differential) enter result for percent Neutrophils, Lymphocytes, Monocytes, Eosinophils and Basophils. - BF Macrophage and Mesothelial cells (if applicable): Report as Occasional, Few, Moderate, and Many - Comments: Report unidentifiable and immature cells under comments. Note: If specimen is clotted, report RBC and WBC count as “Not Measured” and add in comment “Clotted Specimen.” Differential can be reported with comment e.g. “Results may be inaccurate, specimen is clotted.” |
| 2. | Results are manually entered in Cerner under the Accession Result Entry (ARE) mode. For detailed instructions of entering results, please refer to Laboratory Informatics – Cerner Genlab Policies & Procedures Manual, “Resulting in Cerner GenLab: Manual Entry” LIS.SCPMG.041 document. |

| ##### REFERENCE RANGE | Appearance: clear White Blood Cells: 0-8 |
| --- | --- |

| ##### REFERENCE | Todd-Sanford. Clinical Diagnosis of Laboratory Methods 18th Edition. Davidson and Henry, M.D. Brays Clinical Laboratory Methods 7th Edition. p. 512 |
| --- | --- |

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Name of responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| Minor | Added synovial fluid as a body fluid, added gross exam block, updated the resulting count portion and reference. | Diana Davodi, 6/22/10 |  |  |  |
| Major | 1) Updated format, added the materials/reagents section. 2) Revised gross examination procedure to reflect current process. 3) Updated the result reporting section to reflect the changes for new LIS, Cerner. | Julius Salomon 5/12/14 |  |  |  |

Imp. =Implemented