| ##### INTENDED USE | M-TP reagent, when used in conjunction with the UniCel® DxC 800 System and SYNCHRON® Systems Microprotein Calibrator, is intended for the quantitative determination of total Microprotein human urine and cerebrospinal fluid (CSF). An increase in spinal fluid protein is seen in a variety of disease states. Among them are meningitis, polyneuritis, and some tumors. Increases in urinary proteins are associated with a number of conditions among them are nephrosis hypergammaglobulinemia, pregnancy, and destructive lesions of the kidney. |
| --- | --- |

| ##### METHODOLOGY | M-TP reagent is used to measure the protein concentration by a timed endpoint method. Protein in the sample reacts with the pyrogallol red (PR) and molybdate (Mo) to form a purple color complex that has a maximum absorbance at 600 nanometers. |
| --- | --- |

| ##### SAMPLE | It is recommended that urine assays be performed within 2 hours of collection. For timed specimens, the collection container should be kept in the refrigerator or on ice during the timed period. Preservatives are not recommended. - CSF specimens should be centrifuged and analyzed without delay. Specimens may be refrigerated or frozen for 7 to 10 days for repeat determinations. |
| --- | --- |

| ##### REAGENTS | Beckman Microprotein Reagent Cartridge stored at +2°C to +8°C - No preparation is required. - Onboard stability is 30 days. |
| --- | --- |

| ##### CALIBRATION | - Calibrators: Synchron Systems Microprotein Calibrator stored at -15°C to -20°C - No preparation is required. - The system must have a valid calibration in memory before controls or patient samples can be run. Under typical operating conditions the mTP assay must be calibrated every 14 days or with each new reagent cartridge loaded and also with certain parts replacements or maintenance procedures. (See DXC 800 Instructions for Use Manual for complete calibration instructions.) |
| --- | --- |

| ##### CONTROLS | Biorad CSF Controls 1 and 2 and Biorad Urine Chemistry Controls 1 and 2 run daily. In addition, these controls should be run with each new calibration, with each new reagent bottle loaded, and after specific maintenance or troubleshooting procedures as detailed in the appropriate system manual. |
| --- | --- |

| ##### SAFETY | All reagents and controls should be handled as though capable of transmitting infectious diseases. Wear appropriate personal protective equipment when running patient samples or performing scheduled maintenance. Refer to: Policy and Procedures Safety Manual Infection Control and Procedures 11-030-01 |
| --- | --- |

| ##### PROCEDURE | Program samples for analysis as directed in 04-010-01. Primary tube samples must have 264 μL of samples for testing. Use Beckman Primary Tube Sample Template (see Chemistry Information Manual I – Quick Reference Guide) for visual assessment. Specimens placed in a 0.5 mL sample cup require at least 10 μL (includes 40 μL for dead space). FOR PROCESSING 24 HOUR URINE SPECIMENS: - Verify that the Urine Total Volume has been resulted in Cerner - Load urine sample with barcode onto DxC. - Result will come across in Instrument Queue with all fields completed. - Verify results. |
| --- | --- |

| REFERENCE RANGE | SAMPLE TYPE | CONVENTIONAL UNITS |
| --- | --- | --- |
|  | CSF | 15 – 45 mg/dL |
|  | Urine (random) | < 10 mg/dL |
|  | Urine (24 hour) | 50 – 100 mg/24 hrs |

| ##### ANALYTICAL RANGE | 6 - 300 mg/dL (CSF) 6-150 mg/dL (urine) CSF and urine samples with concentrations exceeding the high end of the analytical range should be diluted with normal saline and reanalyzed. |
| --- | --- |

| ##### REFERENCES | Beckman Synchron Clinical Systems Chemistry Information Manual, 2007 http://kpnet.kp.org:81/california/scpmg/labnet/ |
| --- | --- |

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| Major | Added random and timed urine as acceptable specimens for testing. | Cindy Schwartz 2/10/12 |  |  |  |
| Minor | Changed resulting method for 24 hour urine total protein | Cindy Schwartz 9/18/14 |  |  |  |