#### STA-R STARTUP AND SPECIMEN PROCESSING for PROTIME, PTT, FIBRINOGEN, & QUANTITATIVE D-DIMER

| ##### PRINCIPLE | The STA–R® Evolution is a fully automated coagulation instrument that uses an electromagnetic viscosity detection system. The oscillation of a steel ball within the cuvette with the reagent and diluted plasma is monitored by the STA–R® Evolution. The viscosity increases as a result of the coagulation, the oscillation amplitude of the ball swing decreases. An algorithm uses these variations in oscillation amplitude to determine the clotting time in seconds. |
| --- | --- |
| ##### WORKPLACE SAFETY | Laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures. Refer to Policy and Procedures Safety Manual Infection Control 11-085-01. |

| ##### DAILY STARTUP | Print and review products available on instrument. From review, plan reconstitution for all analytes needed. All required maintenance must be performed with the frequency recommended by the manufacturer. CLS must sign the maintenance log provided. |  |
| --- | --- | --- |
| ##### CONTROL & REAGENT PREPARATION |  |  |

| ##### CONTROLS: |  |
| --- | --- |

| STAGO COAG & LIATEST CONTROLS | Reconstitute each vial of control with exactly 1 mL of reagent grade water (See Laboratory P&P 01-130-01). Allow the reconstituted material to stand at room temperature (18-25°C) for 30 minutes. Then, swirl vial gently before use. All reagents should be dated and initialed before placing on instrument. Reagents are stable for 8 hours on the analyzer (15-20°C). |
| --- | --- |

| ##### REAGENTS: |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | APTT | - Reagent 1: PTT-Automate: Reconstitute each vial with 5.0 ml of reagent grade water (See Laboratory P&P 01-130-01). Let sit 30 minutes at room temperature. Swirl gently. Remove the stopper and replace the perforated plastic cap. Place in the R1 section of the drawer. Reconstituted stability on the STA-R is 24 hours. - Reagent 2: 0.025 M CaCl2: Ready to use. Remove the rubber stopper and cap from the vial. Place in the R2 section of the drawer. Stability is 120 hours on the STA-R. |  |  |
|  | PT | - Reagent 1: Neoplastine CI+: Transfer the entire contents of one vial of Reagent 2 into one vial of Reagent 1 of the same lot. Use a transfer pipette to aspirate all the remaining reagent. Let sit 30 minutes at room temperature. Swirl gently. Add a stir bar and maxi-reducer to the vial and place perforated plastic cap on the vial. Make certain the stir bar is not caught underneath the reducer. Place the vial into a stirring position in the reagent drawer (R2). Reconstituted stability on the STA - R is 48 hours. - Reagent 2: 10 ml solvent. Ready to use. |  |  |
|  | FIB | - Reagent 1: STA-Fibrinogen: Reconstitute each vial with 5.0 ml of reagent grade water (See Laboratory P&P 01-130-01). Replace the rubber stopper and perforated plastic cap on the vial. Let sit 30 minutes at room temperature. Swirl gently. Remove the rubber stopper and place the perforated plastic cap on the bottle. Scan reagent and place the reagent into the reagent drawer (R2). Reconstituted stability on the STA - R is 120 hours. - Reagent 2: Owren-Koller buffer: Ready to use buffer. Used by the STA -R to perform dilutions of controls and patients’ plasmas. Place the reagent into the reagent drawer (R0). Stability is 120 hours on the STA – R. |  |  |
|  | D-DIMER | - Reagent 1: Liatest D-Di Buffer: Ready to use. Allow the reagent to stand at room temperature for 15 minutes. Mix gently without creating bubbles. Then, remove rubber stopper, insert a STA – R mini Reducer, and place the perforated cap on the vial. Place the reagent into the reagent drawer (R1). Reconstituted stability on the STA – R is 15 days. - Reagent 2: Latex Liatest D-Di: Ready to use. Allow reagent to stand at room temperature for 15 minutes. Mix gently without creating bubbles. Then remove stopper. Insert a STA – R mini Reducer, and place the perforated cap on the vial. Place the reagent into the reagent drawer, (R2). Reconstituted stability on the STA – R is 15 days. |  |  |

| ##### LOAD REAGENTS | Loading Stago (controls and reagents) products |
| --- | --- |

| ##### USING BARCODED PRODUCTS |  |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | - Click on the product icon. Click the “OPEN” tab. After drawer opens - tab reads “Close Drawer”. |
| 2. | - Click the ''List of Products" tab. |
| 3. | - Product Loading Screen will be displayed. |
| 4. | - Press “Position” button, then “type” button. This arranges products for easy assessment. |
| 5. | - Scan bar-coded product. Volume and stability times must be edited as needed. For Micro cups click on Microcup icon or the key F8. |
| 6. | - Place the vial in the appropriate position. R0 - Controls, calibrators, and diluents R1 - Reagents added before incubation. R2 - Start reagents (Thromboplastin, CaCl2, FBG Reagent) Note: Desorb U needed in all positions (RO, R1, & R2) |
| 7. | If a new lot is detected, the following question is displayed “Do you wish to read the calibration parameters using the barcode reader?" To proceed directly with the barcode reading, click OK, pass the sheet in front of the barcode reader and then click Confirm. |
| 8. | Load all products necessary. |
| 9. | Click on “Close Drawer”. |

| ##### USING MANUAL IDENTIFICATION |  |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Open the Product Drawer. |
| 2. | Key in the product ID as defined in Test Setup, then click ENTER. |
| 3. | Click on the micro cup icon or the F8 key if loading micro cup. |
| 4. | Key in lot number. |
| 5. | Volume and stability must be edited as needed. |
| 6. | Place vial in appropriate position. |
| 7. | Load all products necessary following steps 2-6 above. |
| 8. | Close drawer. Note: always run controls after loading new reagent. |

| ##### RUNNING CONTROLS | Controls must be run at least once per shift. When Quality Control tolerance limits are exceeded, corrective actions must be performed and recorded before analyzing patient samples. Controls must run in the same manner as the patient samples. |
| --- | --- |

| PT/PTT/FIB | Coagulation Control N and ABN |
| --- | --- |
| D-Dimer | Liatest Control N and P |
| Heparin | Heparin Control Level 1 and 2 |

| Step | Action |
| --- | --- |
| 1. | Running 1 Test by level: - Click the Quality Control icon - Locate the test for which a control must be run - Click the test abbreviation. Background turns to blue - Click the Control Level # tab (i.e. Control Level Tab 1) - Click Run Selection icon (pointing finger). - Confirm by clicking OK - Repeat for all levels to be run |
| 2. | Running multiple tests: - Click the Quality Control icon - Click on the uppermost test abbreviation of test to run. Background turns to blue - Hold down the ‘Ctrl” (left) key on the mouse pad, pull mouse arrow across the other tests to run all. Backgrounds turn blue. Lift “Ctrl” key. - Click the Selection icon (pointing finger). - Confirm by clicking OK |

| ##### Validating, Deleting, and Rerunning QC |  |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Click the Quality Control icon flashing on the bottom of screen. |
| 2. | Locate the test for which a quality control must be validated |
| 3. | Double click the test that must be validated. |
| 4. | A box opens up-select validate. |
| 5. | The alarm will stop beeping. |
| 6. | You may double click on the box again if you wish to rerun or delete that control result. |

| ##### RUNNING PATIENTS | Follow the steps below to process patient samples. At the beginning of each shift, whenever preventative maintenance is done, or new reagent is put in instrument, run controls from the reagent holder work list. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Check label to make sure barcode is readable and tube can rotate freely in rack. |
| 2. | Place tubes in rack. |
| 3. | Slide rack(s) on to tray. |
| 4. | Place tray on STA-R loading station: make sure it is seated properly. Make sure automatic inquiry icon is pictured at the bottom of the test panel screen. - Racks will be loaded and barcodes read. - Tubes with no barcodes or barcodes that can not be read are assigned an ID # (NL/# of rack/position in rack) NOTE: All fibrinogens are run in duplicate and must agree within 10% of each other. |
| 5. | Ordering tests can be done by the following: - Using Downloading (orders will download from Cerner and tests will run automatically). - Using Automatic Profile (same tests done on all samples) - Tagging IDS and ordering individual tests on all tagged IDS. - Individual tests can be ordered by double clicking in the box of the test needed. NOTE: For detailed instruction refer to Operator's Manual - pg.19-23. |

| Loading STAT rack |  |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Click on STAT icon |
| 2. | When 'READY" screen is displayed, push rack against white tooth on conveyor belt; belt will take over and push rack into star (barcode side of rack first). |

| ##### Ordering/Rerunning a Test from Test Panel Screen using the Test Box |  |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Double-click the test box for the desired test window. Window will open. |
| 2. | Choose the appropriate selection then click on “OK” |

| Ordering/Rerunning a Test from the Test Panel Screen using Tag Function. |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Step | Action |  |  |
|  | 1. | Tag the appropriate ID or IDs. (Red check mark appears in left hand box under “T”. |  |  |
|  | 2. | On the Test Panel Screen, click on the Test Abbreviation (in the gray column header). A box opens up. |  |  |
|  | 3. | Select Add or Rerun. |  |  |
|  | 4. | Confirm by clicking the OK button. |  |  |
|  | 5. | The selected test has been added/rerun to all tagged samples. |  |  |

| Tagging: There are three ways to tag Identities. They are: |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Step | Action |  |  |  |
|  | 1. | Click on T in gray header column - Click Tag all - All samples loaded on STA-R will be tagged |  |  |  |
|  | 2. | Move cursor to a sample - Click on the box to the left of the ID number (under T) - Clicking on this box to the left will tag or untag the sample |  |  |  |
|  | 3. | Use the Tag icon (Lower right corner of Patient list) - Each click of this icon will tag or untag the sample highlighted by the cursor. The cursor then moves down to the next sample |  |  |  |

| Unloading several/all racks |  |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Click the Patient icon. |
| 2. | Click the loading tab. |
| 3. | Tag racks to be unloaded or tag all. |
| 4. | Click Rack icon. The following message is displayed: Unload the tagged racks? |
| 5. | Confirm by clicking OK. The icon representing the rack shows a rack icon with the word STOP in front. |
| 6. | Release the tray by clicking on the Tray icon. The tray area on the screen turns green. |
| 7. | The tray can then be lifted and removed from Star. |
| 8. | When the tray is removed, the patient’s IDS are removed from the test panel screen and transferred to the archives. |

| ##### RESULTS REPORTING: PT, APTT, FBG & QUANT. D-DIMER | - Results will appear on Cerner identified by the barcode number and patient demographics. - For the current reportable range refer to P&P 01-100-01. - See Lab P&P 05-090-01 for procedure regarding notification of critical values. |
| --- | --- |

| ##### REPORTING GUIDELINES |  |
| --- | --- |

| Analyte | Reportable Range | Exceptions |
| --- | --- | --- |
| PT | 0.5 – 10.0 INR (8-150 sec.) | If INR is >10 (or >150 sec.) or V>V-MAX, no value will crossover to Cerner, report as “>10.0 INR”, this will trigger the critical pop up notification. |
| APTT | 15-200 sec. | If >200 sec. or V>V-Max, no value will crossover to Cerner, report as “>200 sec”, this will trigger the critical pop up notification. |
| FBG | 60-1800 mg/dL | If <60 mg/dL, report as “<60 mg/dL”, this will trigger the critical pop up notification. If >1800 mg/dL, report as “>1800 mg/dL”, this will trigger the critical pop up notification. NOTE: A VMAX result means value is extremely low. A VMIN result means value is extremely high. |
| D-Dimer (DVT, PE) | 270-4000 ng/mL FEU | If V>VMAX, no value will crossover to Cerner. You must check the instrument to see if the result is “V>VMAX” and if so, report as “>4000 ng/mL.” |
| D-Dimer (DIC) | 270-20,000 ng/mL FEU | If V>VMAX, no value will crossover to Cerner. You must check the instrument to see if the result is “V>VMAX” and if so, program instrument to run D-DI x 5 (instrument does automatic dilutions) and if result is >VMAX, report as “>20,000 ng/mL.” |

IMPORTANT: 	DO NOT “Convert Results” to “Freetext”, just type in < or > and the value, it will default to what is in Cerner and trigger the critical notification pop up if applicable.

- REFERENCE           See Coagulation P&P 01-100-01

- RANGE

| ##### THERAPEUTIC RANGE: ##### PT INR |  |  |  |
| --- | --- | --- | --- |
|  | PROTIME/CLINICAL STATE |  |  |
|  | Prophylaxis - venous thromboembolism (high risk surgery) |  |  |
|  | Treatment of initial episodes of DVT or PE |  |  |
|  | Prevention of systemic embolism in-patients with atrial fibrillation (without systemic embolism), valvular heart diseases, tissue heart valves, acute myocardial infarction anterior transmural) and atrial fibrillation with systemic embolism |  |  |
|  | Mechanical prosthetic heart valves and recurrent systemic embolism |  |  |

PREVENTATIVE

MAINTENANCE/

SUPPLY

SUPPLY MANAGEMENT

- Changing the cuvette roll: do this when all the cuvettes are used up in the current roll

| Step | Action |
| --- | --- |
| 1. | Open the upper right door. |
| 2. | Raise the catch holding the roll in place and catch holding the reel in place. |
| 3. | Open the door of the cuvette loader carefully. DO NOT ALLOW TO SLAM OPEN. |
| 4. | Remove the reel and roll by sliding each off its spindle. |
| 5. | Discard the reel and the roll. Note: New cuvette rolls always include a reel which is connected to the roll by a film on which the cuvettes are fixed. |
| 6. | When reading the cuvette roll bar code data, hold the bar code of the new cuvette roll in front of the product drawer bar code reader. If these 2 bar code labels are not detected by the reader, the information from both bar code labels can be entered manually. |
| 7. | A cuvette loading window opens: showing the number of cuvettes and the code\batch of the roll that has just been loaded. |
| 8. | Hold the new cuvette roll in such a way that the roll is on the left and reel is on the right. Lot # is facing you. |
| 9 | Slide the roll along its spindle and the reel along its spindle. Lower the catch holding the reel in place, and the catch holding the roll in place. |
| 10 | Position the film in the cuvette loader so that it is parallel to the loading racks, and close the cuvette loader door. Note: Do not twist the film. |
| 11. | Press the button located beside the reel so as to activate the cuvette feeder. Hold button down steadily until it stops on it’s own. As it moves, check that the cuvettes correctly enter the cuvette loader without twisting sideways and continue to press the feed button until the cuvettes have reached the other side of the loader. The motor stops automatically. |
| 12. | Click the Products icon and then the Cuvette button. The following on screen message is displayed: “Once you have replaced the Cuvette Roll, confirm quantity". Type in 1,000. Hit the green curvy arrows icon. |

|  | Changing the cuvette disposal bag |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Open the front door. |
| 2. | Remove the cuvette bin (on the extreme right). |
| 3. | Lift the plastic bag from the metal container in order to close it. |
| 4. | Remove the tie along the opening of the plastic bag. |
| 5. | Close the plastic bag by forming a knot with the tie. |
| 6. | Dispose of the plastic bag, observing the hospital's regulations. |
| 7. | Take one of the plastic bags located in the box holding the cuvette rolls. |
| 9. | Open the plastic bag and place it in the metal bin. |
| 10. | Ensure that the bag fully spreads out at the bottom of the metal bin. |
| 11. | Lower the plastic bag along the sides of the metal bin. |
| 12. | Put back the metal cuvette bin into the STA-R ensuring that it is down fully in its holding place. There should be an audible click signaling the presence of the bin. Do not disrupt the function following the detection of the presence of the metal cuvette bin. |

Loading the cleaning solution: do this when the instrument alarms to let you know that the waste is full.

|  | Step |  | Action |  |
| --- | --- | --- | --- | --- |
|  | 1. |  | Click the Product icon and STA-Cleaner solution button: a window is displayed with message "Remove and dispose of the Wash Bottle. CAUTION: Biohazard |  |
|  | 2. |  | Open the front door. The bottle of used solution contains potentially infectious material and must be disposed of according to the hospital regulations. |  |
|  | 3. |  | Gently pull out the bottle of used solution shown in blue on the screen (bottle in the center). |  |
|  | 4. |  | Remove the cap of the bottle as soon as possible. Note: A symbol is used at cap level of the used solution bottle and the tubing going into this cap is red in order to indicate that the contents of this bottle is potentially infectious. |  |
|  | 5. |  | Fully remove the bottle of used solution: a new window is displayed with the message “Empty then move the Cleaner Solution bottle." |  |
|  | 6. |  | The empty STA-Cleaner Solution bottle becomes the new waste bottle. |  |
|  | 7. |  | Gently pull out the empty bottle of STA Cleaner Solution (bottle on the extreme left). |  |
|  | 8. |  | Remove the cap from the bottle as soon as possible. Avoid potential contamination by leaving the tube to hang down outside the instrument. |  |
|  | 9. |  | Empty the bottle of STA-Cleaner Solution. |  |
|  | 10. |  | Place the empty bottle at the front of the middle position and put the cap back on the empty bottle. |  |
|  | 11. |  | Push the empty bottle back into place. There should be an audible click signaling the detection of the presence of the bottle: a new window is displayed with the message "Install a new Cleaner Solution.” Do not disrupt the function of the switch allowing the detection of the presence of the bottle of used solution. |  |
|  | 12. |  | Take a new bottle of STA-Cleaner Solution. |  |
|  | 13. |  | Place the new STA-Cleaner Solution in front of the empty space on the extreme left to insert the tube, and then insert tube and connect cap to bottle. |  |
|  | 14. |  | Push the empty bottle to the back of its recess. There should be an audible click signaling the detection of the presence of the bottle: a new window is displayed with the message "Indicate the volume." |  |
|  | 15. |  | Confirm the volume shown (2500 ml) by clicking Validate if a new bottle is used. If not, enter the remaining volume, and confirm by clicking Validate. |  |
|  | 16. |  | Close the right front door. |  |

- PREVENTIVE MAINTENANCE

Refer to STA-R Evolution Reference Manual section 4 for detailed instructions on performing each task.

| Frequency | Tasks |
| --- | --- |
| Daily | - Check reagent arm #3 temperature - Check measuring block temperature - Check reagent drawer - Record pipetting % of syringes #1,2 & 3 - Cap- 10 -Piercer Daily Maintenance - Clean the Touch screen |
| Weekly | - Clean the main filter - Clean the washing wells with 10% bleach - Clean rising well N0. 3 and needle N0.3 - Cleaning the product drawer - Clean the measurement plate - Clean the suction tip - Clean the conveyor belt - Shuttle cleaning |
| Monthly | - Replacing the syringe teflon tips |
| Quarterly | - Replacing air filter - Replacing the barcode reader wheel |

| ##### REFERENCE | DIAGNOSTICA STAGO Reference Manual October2005, January 2006 |
| --- | --- |

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Name of responsible person/date | Med. Dir. reviewed/ date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| Major | 1) Added reference to reagent grade water source. 2) Remove any reference to LMWH. 3) Replaced any reference to LMS with Cerner. 4) Remove any reference to repeating any critical results. 5) Remove any reference to D-dimer results reporting protocol, P&P Office Procedures 04-555. | Julius Salomon 08/05/13 |  |  |  |
| Major | Page 7 – Removed “Coagulation Repeats/Critical Values” section. And added See Lab P&P 05-090-01 on Results Reporting section. Page 7 - Revised reporting procedure for results that’s either < or > than reportable range to reflect current workflow in Cerner. Also revised reportable ranges to correctly reflect Cerner (LIS) ranges. | Julius Salomon 11/15/13 |  |  |  |
| Major | Page 7 1) added INR for PT reporting 2) added D-Dimer for DIC reporting | Julius Salomon 12/7/15 |  |  |  |

Imp. =Implemented