| Dept. | CHEMISTRY | Policy #: |  | 04-002 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Policy/ Procedure Title | DXC800 STARTUP, SHUTDOWN AND REBOOT | Effective Date: |  | 05/13/16 |  |  |
|  |  | Revised Date: |  |  |  |  |
|  |  | Authorized Approval: |  | Cindy Schwartz, CLS, Area Lab Manager Jana Pindur M.D., Laboratory Director |  |  |
| Personnel Covered: | All Chemistry Staff |  |  |  |  |  |
|  |  | Page: | 1 |  | of | 2 |

Purpose

To define proper procedures for power down, start up and reboot of the DxC 800 system.  This is necessary for conditions such as unexpected power failure, computer down time and trouble shooting.

Procedure

## Power Down System

- With the system in standby.

- Select Instrument commands from Menu Bar.

- Select <4> Shutdown.

- When the following message appears on the screen “Shutdown in progress” select [Continue].

- When the following messages appear on the screen “The console is now ready to be powered off”.

- Power down PC using On/Off switch on PC.

- Power down Analyzer using Main Switch behind lower right hand door.

Power Up System

To power up and resume operation.

- Power on Analyzer using Main Switch.

- Power on Computer.

System will automatically Reboot ~ 10 min.

When the main screen appears and status display highlights: Standby resume normal operations.

Reset/Reboot System

- With the system in standby.

- Select Instrument commands from Menu Bar.

- Select <4> Shutdown.

- Select <OK>.

- When the following message appears on the screen “Shutdown in progress” select [Reboot].  System will automatically REBOOT approximately 10 min.  When instrument status displayed on screen reads “Standby” resume normal operations.

References

Unicel DXC Synchron Clinical Systems onboard IFU

www.beckmancoulter.com

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP - describe | Signature responsible person/date | Med. Director Authorized Reviewed/Date | Lab Manager Authorized Reviewed/Date | Date change Implemented |
| --- | --- | --- | --- | --- | --- |
| Major | New | Cindy Schwartz 5/13/16 |  |  |  |